| Journal Name: | Asian Journal of Medicine and Health |
| --- | --- |
| Manuscript Number: | Ms_AJMAH_136896 |
| Title of the Manuscript: | RADIATION DOSE AND IMAGE QUALITY OF ROUTINE CT EXAMINATIONS |
| Type of the Article | Original Research Article |

| ## PART 1: Comments |  |  |
| --- | --- | --- |
|  | ## Reviewer’s comment Artificial Intelligence (AI) generated or assisted review comments are strictly prohibited during peer review. | Author’s Feedback (It is mandatory that authors should write his/her feedback here) |
| Please write a few sentences regarding the importance of this manuscript for the scientific community. A minimum of 3-4 sentences may be required for this part. | The topic of this manuscript is relevant and timely, considering the ongoing emphasis on optimizing radiation dose while maintaining diagnostic image quality in computed tomography (CT), including comparison of anatomical regions and gender. The investigation of the relationship between radiation dose, image noise, and signal-to-noise ratio (SNR), particularly in relation to body mass index (BMI), is scientifically justified and contributes to the literature on personalized imaging protocols. However, despite the study's premise is valid, the technical parameters assessed (signal, noise, and SNR) are well-established and widely studied. To enhance the scientific contribution, consider including more advanced or complementary metrics such as contrast-to-noise ratio (CNR), task-based assessments, or qualitative evaluations by expert radiologists. The study relies on CTDIvol and DLP to describe radiation dose; however, these are scanner output metrics and do not account for patient size. Especially in the thorax and abdomen regions—where tissue attenuation varies considerably with BMI—the use of Size-Specific Dose Estimates (SSDE) is highly recommended. SSDE provides a more accurate, patient-specific dose estimate and is advocated by the AAPM (Report No. 204). Including SSDE values would significantly strengthen the validity of the dose assessments in this study. | ## Thank you for this critical and healthy review, the focus of CNR is not covered in this article. To further enhance the scientific contribution as suggested, the figure of merit (FOM) has been included as a complimentary metric. The SSDE has also been included as recommended to strengthen the validity of dose assessments. |
| Is the title of the article suitable? (If not please suggest an alternative title) | It is general and ambiguous. Something like: "Impact of Body Mass Index on Image Quality and Radiation Dose in Routine CT Scans: Emphasis on Dose Optimization Strategies", or “Impact of Body Mass Index & Sex on Image Quality and Radiation Dose in Routine CT Scans” | ## The title of the article has been changed to the suggested title (Impact of Body Mass Index and Sex on Image Quality and Radiation Dose in Routine CT scan) |
| ## Is the abstract of the article comprehensive? Do you suggest the addition (or deletion) of some points in this section? Please write your suggestions here. | No, it is OK | ## Thank you |
| ## Is the manuscript scientifically, correct? Please write here. | Yes. | ## Thank you |
| Are the references sufficient and recent? If you have suggestions of additional references, please mention them in the review form. | A substantial proportion of the cited literature is more than 5 or even 10 years old. Given the rapid technological advancements in CT imaging—especially in the fields of iterative reconstruction, imagen quality assessment, artificial intelligence-based dose optimization, and advanced dose modulation techniques—it is strongly recommended to incorporate more recent references (from the last 3–5 years). Examples: Steiniger B, Fiebich M, Grimm M-O, Malouhi A, Reichenbach JR, Scheithauer Marcel, et al. PAE planning: Radiation exposure and image quality of CT and CBCT. Eur J Radiol 2024;172:1–8. Ukaji NF, Ohagwu CC, Promisegolodom M. Effects of Variations in Imaging Parameters on Image Quality of Non Contrast Computed Tomography Scans of Brain: A Cross-sectional Study. Journal of Clinical and Diagnostic Research 2021;15:13–7 Wong YM, Ong CC, Liang CR, Tan CA, Teo LLS. Image quality, contrast enhancement and radiation dose of electrocardiograph- versus non-electrocardiograph-triggered computed tomography angiography of the aorta. Singapore Med J 2024;65. Chatzaraki V, Kubik-Huch RA, Thali M, Niemann T. Quantifying image quality in chest computed tomography angiography: Evaluation of different contrast-to-noise ratio measurement methods. Acta Radiol 2021;63:1353–62. https://doi.org/10.1177/02841851211041813. López A, Rodríguez DD, Mendoza RN, Garrido J, Martín JM, Del Pozo A. CT image quality metrics implementation during thorax SPECT/CT clinic studies. Revista Brasileira de Física Médica 2023;17:724. https://doi.org/10.29384/rbfm.2023.v17.19849001724. Ohashi K, Nagatani Y, Yoshigoe M, Iwai K, Tsuchiya K, Hino A, et al. Applicability Evaluation of Full-Reference Image Quality Assessment Methods for Computed Tomography Images. J Digit Imaging 2023;36:2623–34. https://doi.org/10.1007/s10278-023-00875-0. Lee W, Cho E, Kim W, Choi H, Beck KS, Yoon HJ, et al. No-reference perceptual CT image quality assessment based on a self-supervised learning framework. Mach Learn Sci Technol 2022;3:045033. https://doi.org/10.1088/2632-2153/aca87d. | ## The references have been updated to include only literatures from the last 5 years as suggested. Most of the suggested articles here have been incorporated into the reference |
| ## Is the language/English quality of the article suitable for scholarly communications? | OK | We appreciate this comment |
| ## Optional/General comments | While the study provides a quantitative assessment of image quality parameters, it would benefit from a discussion of the clinical implications of these findings. For instance, how does the observed variation in SNR impact diagnostic performance in evaluating lesions or anatomical structures in patients with high BMI? | acknowledged |

| PART 2: |  |  |
| --- | --- | --- |
|  | ## Reviewer’s comment | Author’s Feedback (It is mandatory that authors should write his/her feedback here) |
| Are there ethical issues in this manuscript? | (If yes, Kindly please write down the ethical issues here in detail) Not included in the manuscript | No ethical issues |