QUANTITATIVE ASSESSMENT AND COMPENSATION OF TEMPERATURE-INDUCED ERRORS IN LOAD CELL MEASUREMENTS USING THERMAL DRIFT AND HYSTERESIS ANALYSIS.

ABSTRACT

Accurate weight measurement is a fundamental requirement in scientific experiments, industrial process, and precision-based applications. However, load cell-based weighing systems are highly sensitive to environmental conditions, particularly temperature fluctuations, which introduce systematic errors and non-linear variations in the analog-to-digital convertor (ADC) output. The study investigates the influence of temperature on the performance of a 10 kg load cell while measuring weights in the lower range of 200 g to 1000 g across 15–35 deg. C under heating and cooling cycles.. The experimental setup consists of ESP 32 microcontroller, HX711 power amplifier, DHT22 temperature sensor. The weight measurement system being IoT enabled with ThingSpeak IoT platform, the raw ADC counts, temperature and weight data are time stamped to cloud dashboard for continuous logging, remote monitoring. Before correction, the maximum percentage errors rose to 2.39% at 200 g, and to 1.95%, 1.38%, 1.54%, and 1.42% at 400 g, 600 g, 800 g, and 1000 g, respectively, at the upper end of the thermal range. After applying a temperature-aware calibration, errors were substantially reduced and stabilized: peak post-correction errors were 0.52% (200 g) and <0.39% for 400–1000 g (e.g., 0.39% at 400 g, 0.31% at 600 g, 0.28% at 800 g, 0.27% at 1000 g at 35 deg. C), with typical values clustering around 0.2–0.3% across mid-range temperatures. By coupling the compensation pipeline with an IoT architecture, the platform supports real-time analytics, trend visualization, and over-the-air refinement of calibration coefficients, enabling robust field deployment for compensation of temperature variation.

Keywords: ADC, ESP 32, HX711 power amplifier, DHT22 temperature sensor, IoT, ThingSpeak.

INTRODUCTION

A piece of machined metal called a transducer or load sensor is known as load cell. It is fabricated in such a way that it is bound to bend with mechanical force and therefore convert the mechanical force into electronic signal. The signal conversion happens with the aid of strain gauge embedded into the metal. The load cell, also known as force sensor has several advantages such as high precision, high sensitivity, high linearity, small size and mature manufacturing technology(Zhuang et al., 2023). Measurement of perfect weight is most vital requirement. It is very important to measure precisely and store the measured value of weight. The mechanical weighing machines are replaced by the electronic weighing machine as the electronic weighing machine has upper hand of higher accuracy and higher degree of reliability. Compared to other weighing machine load cell based has advantages like high dependability, facility of online processing through computer, print out facility and attachment of parallel display(Das et al., n.d.-a).

One of the first use of bonded resistance wire strain gauge following its discovery in the early 1940s was in the development of an accurate and reliable load cell or force transducer. The operating principle suggests that if a wire is bonded to a spring element in such a manner that its cross section varies as the spring element is strained then it is possible to establish a relationship between the electrical resistance of the wire and the force causing the deformation of the spring element(Ajovalasit, 2011). Strain gauge load cells are designed for operation within specific capacity ratings. Excessive overloads may result in loss of accuracy or failure. In general, the load cells should not be subjected to more than 125% of their rated capacity. Typical types of strain gauges used are semiconductor strain gauge, nuclear radiation sensor, inductive sensing, variable reluctance sensing. In semiconductor strain gauge the piezoelectric characteristics of germanium and silicon materials are harnessed. They are most commonly used in the construction of load cells as they possess very high strain sensitivity, substantial nonlinearity, and temperature effects on strain sensitivity and terminal resistances are also relatively high(Ling et al., 2024). Inductive and reluctance load cells incorporate the two basic sensing principles i.e. the motion of ferromagnetic core (inductive) or a coil assembly (reluctance) is converted to a voltage signal directly proportional to the displacement. Various force sensing element converts the applied force or weight to a displacement to which the sensing element is coupled. This type of transducer furnishes relatively high output signal and moderate to high accuracies and cover a broad range of measuring capacities(Chen & Conn, 2020). For measuring heavy weights load cells built in with microprocessors are used. An example of the load cells that are provided with built-in microprocessor are the Toledo DigiTOL powercells. They are able to withstand the adverse weather conditions in outdoor sites. Each powercell contains a microprocessor that has been programmed to automatically recognize and correct error caused by external influences or loads moving on the scale base while weighing is in progress. The capacity of the weighbridge is approximately 500 tons(Hastawan et al., 2021).

The measuring device which is load cell and the entire framework is prone to environmental forces which can produce errors in the load cell signal. The most significant force or input modifier is high ambient temperature. But it is not possible to keep up a controlled environment. Changes in surrounding temperature cause load cell sensitivity to drift. The change on the load cell to zero load causes the whole cell’s output range to shift(Abhiram, 2022). During the production and use of load cells, many performance compensations are required like hysteresis compensation, creep compensation, zero compensation and sensitivity improvement. When the external temperature changes, the thermal expansion and contraction effects and residual stress affect the performance of the load cell(Luo et al., 2012). According to OIML R60 regulations, as long as the performance compensation at -10℃, 20℃, and 40℃ is within the acceptable range, the performance compensation of the load cell is considered to be up to standard(Bétaille & Jacob, n.d.).

Accurate weight measurement is critical in diverse applications ranging from industrial production and laboratory research to commercial trade. Load cells, which convert applied force into an electrical signal, are widely used in digital weighing systems due to their high sensitivity and reliability. However, their performance is significantly influenced by environmental factors, with temperature variation being one of the most prominent sources of error. Changes in ambient temperature affect both the mechanical properties of the load cell and the electrical resistance of its strain gauges, leading to deviations between actual and measured weight. In precision applications, even minor errors can accumulate into substantial inaccuracies, impacting decision-making and operational efficiency. This study focuses on quantifying and correcting temperature-induced errors in a custom-built weight measurement system using a 10 kg load cell. Known standard weights of 200 g, 400 g, 600 g, 800 g, and 1000 g were measured across a temperature range of 20 °C to 40 °C to identify error patterns. Subsequent experiments from 15 °C to 35 °C during both heating and cooling cycles were conducted to validate temperature compensation methods. By developing empirical correction equations from raw ADC readings, this research aims to enhance measurement accuracy, enabling the use of low-cost weighing systems in precision-demanding environments.

LITERATURE REVIEW

The study investigates the effect of temperature variation on the accuracy of a single-point load cell and develops a simulation model using TRNSYS 16.0 to estimate temperature-induced errors for Zaria, Kaduna State, Nigeria(Edet et al., 2022). The objectives included measuring real-world load cell readings across a range of temperatures, comparing them to true values to quantify errors, simulating load cell performance with Zaria’s temperature data, and creating predictive models for uncompensated systems. A single-point load cell was tested under controlled conditions using standard weights, with measurements taken at various ambient temperatures reflective of local climate conditions. The difference between measured and actual weights was calculated to determine errors, and historical temperature data from Zaria was used in TRNSYS 16.0 to simulate long-term error patterns using load cell response equations. Results showed significant drift in readings with temperature changes, exhibiting a non-linear pattern with both positive and negative deviations depending on temperature direction, and certain ranges producing larger spikes due to strain gauge material characteristics. The study concluded that temperature variations greatly affect load cell accuracy, making uncompensated systems unreliable in areas with large temperature swings. The close match between simulation results and experimental data validated TRNSYS 16.0 as a reliable predictive tool, and temperature compensation—through hardware or software—was recommended to improve accuracy in practical applications.

The study focuses on designing an optical weight measurement system using Fiber Bragg Gratings (FBGs) and enhancing its accuracy through machine learning techniques(Pal et al., 2022). The key objectives include implementing an optical load cell with FBG sensors embedded in a cantilever structure, developing signal processing algorithms to translate FBG wavelength shifts into weight readings, applying machine learning models such as support vector regression and neural networks to improve accuracy, and analyzing performance in terms of sensitivity, signal-to-noise ratio (SNR), and error characteristics. The methodology involved fabricating a cantilever-type optical load cell integrated with FBG sensors, using a tunable laser and edge-filter detection (EFD) with Direct-Intensity Modulation (D-IM) to measure optical power changes induced by weight-related strain. Standard weights, ranging up to several kilograms, were applied, and the resulting sensor outputs were recorded. These datasets were then used to train machine learning models to predict actual weight from raw FBG intensity or wavelength data. Experimental observations showed that the D-IM approach doubled the sensor sensitivity (≈196.7 mV/kg compared to 87.5 mV/kg) and achieved a high SNR (≈63 dB). While conventional EFD methods exhibited an error of around ±40 g, D-IM reduced this error and improved resolution to approximately 10 g. Furthermore, machine learning calibration further minimized measurement errors and enhanced the system’s capability to compensate for drift and nonlinearities. The study concludes that integrating D-IM with FBG-based optical load cells greatly improves sensitivity, SNR, and resolution, while the addition of machine learning provides further accuracy gains. This combination of lightweight optical sensing technology and intelligent algorithms offers a viable and precise alternative to traditional load cells for industrial applications.

The study focuses on developing a digital device capable of measuring human height, body weight, and body temperature by integrating three different sensors—an ultrasonic sensor for height measurement, a load cell for weight measurement, and an MLX90614 infrared sensor for temperature detection—into a single unit controlled by an Arduino Nano(Saputra et al., 2022). The objectives were to combine these sensors into a cohesive system, design appropriate signal-conditioning circuits for each, process the sensor outputs through a microcontroller, and display the results in a clear, unit-calibrated format on an LCD screen. The hardware setup included an Arduino Nano, ultrasonic proximity sensor, load cell with HX711 amplifier, MLX90614 infrared temperature sensor, LM2596 voltage regulator, and a 16×2 I2C LCD. In operation, the device collects raw readings from each sensor—distance for height, strain-induced voltage for weight, and infrared radiation for temperature—conditions the signals into analog voltages, digitizes them through the Arduino, and processes them to obtain accurate measurement values. These results are then displayed on the LCD, with options for unit selection through the interface. Although detailed numerical accuracy results were not presented, the prototype was reported to function successfully, delivering reliable and accurate digital measurements for all three parameters. The system effectively converts physical phenomena into digital readings, demonstrating consistent performance through ultrasonic echo detection, load cell pressure sensing, and infrared measurement. The study concludes that the developed device serves as a functional and practical prototype, offering a fast, user-friendly, and cost-effective solution for automated body measurement applications using open-source hardware and readily available components.

The paper investigates the internal temperature gradients that develop within a load cell during operation, primarily due to self-heating from strain gauges and PCB electronics, and aims to enhance temperature compensation by optimizing the placement of compensation resistors to minimize errors caused by unstable internal temperature fields(Zhuang et al., 2024). Using steady-state thermal modules in ANSYS/Workbench, the authors simulated heat generation and distribution within the load cell to identify resistor positions that would reduce internal temperature differentials. Based on simulation outcomes, the compensation resistor was repositioned to a more thermally homogeneous region, ensuring that the modification did not interfere with the mechanical strain response of the elastomer. Experimental validation was then conducted by comparing the performance of the modified load cell to the original design. The thermal simulations revealed notable internal temperature gradients, particularly between the strain gauges and the compensation resistors. Relocating the resistor significantly reduced these gradients in the simulation environment, and the experimental tests confirmed that this change improved temperature compensation performance, reducing errors caused by unstable temperature fields. The study concludes that optimal placement of temperature compensation resistors within the load cell’s internal structure can substantially improve measurement accuracy under varying temperature conditions, without increasing hardware costs or compromising strain sensitivity. This resistor placement strategy offers a practical and cost-effective enhancement to conventional load cell temperature compensation methods.

The paper presents the development of an integrated IoT-enabled digital scale system for sheep farming that combines weight measurement, body temperature sensing, and RFID-based animal identification, all controlled by a microcontroller(Ulfa & Isnianto, 2023). The primary aim is to create a system capable of accurately measuring both sheep weight and body temperature while enabling automated animal tracking and real-time data transmission to a remote server or user interface. The designed system consists of a scale frame powered by either a 12 V battery or a 24 V power supply, with a component box housing the load cell module, MLX90614 infrared temperature sensor, Arduino Mega 2560 microcontroller with built-in Wi-Fi (ESP8266), LCD interface, RFID reader, and thermal gun. Each sheep is equipped with a unique RFID tag, which is detected when the animal enters the scale frame, displaying its ID on the LCD. Simultaneously, the load cell measures weight and the thermal gun—positioned toward the sheep’s rectum—captures body temperature, with all measurements sent to a server in real time. Field testing demonstrated reliable RFID detection within a 0–75 cm range, temperature measurement accuracy between 99.69% and 99.97%, and weight measurement accuracy between 98.28% and 99.92% compared to conventional scales. The results confirm that the system delivers high accuracy and operational reliability, providing an effective solution for modernizing sheep farming through automation and remote monitoring.

The study aims to understand how different drop-weight test configurations influence the accuracy of measured impact forces and the resulting force profiles when testing reinforced concrete (RC) beams(Li et al., 2020). Specifically, it seeks to investigate how load cell placement and the mass distribution of the drop weight affect measurement accuracy, to identify configurations that most accurately reproduce the true impact force acting on the beam, and to examine how the mass ratios between drop weight components and between the drop weight and beam influence the force profile. The researchers conducted numerical simulations comparing various drop-weight setups, focusing on two main configurations: embedding the load cell within the drop weight and mounting it between the drop weight and the beam. They analyzed the effects of the mass ratio of the drop weight to its head (αₙ) and the mass ratio between the drop weight and beam on measurement accuracy and impact behavior. Results showed that embedding the load cell in the drop weight can cause deviations between measured and actual contact forces, especially when αₙ exceeds 20, while placing the load cell between the drop weight and the beam alters local contact stiffness and changes the impact force profile. Furthermore, the mass ratio between drop weight and beam (0.5–4.0 range) significantly affects the relative velocity after initial impact, influencing the force-time response. The study concludes that drop-weight test configuration—including load cell placement and mass ratios—plays a critical role in ensuring measurement accuracy and profile consistency. To achieve reliable results, αₙ should remain below 20 when embedding the load cell, and both placement and mass ratios must be carefully optimized. This provides practical guidelines for improving accuracy and comparability in drop-weight impact testing of RC beams.

The study focuses on designing and developing an affordable and accurate digital weight measurement system using an industrial-grade Gefran strain-gauge load cell integrated with a microcontroller platform for real-time display(Shinde, 2025). The primary objectives include building a weight measurement system that utilizes a Gefran strain gauge load cell, interfacing it with an AD620 instrumentation amplifier and Arduino Uno, supplying a stable excitation voltage to the load cell, implementing calibration and software logic for precise real-time readings, and validating the system’s performance up to 200 g with attention to linearity, repeatability, and stability. The methodology involved constructing the system with an Arduino Uno microcontroller, AD620 instrumentation amplifier, and a 16×2 LCD for display. A regulated 10 V excitation voltage was supplied to the load cell via a 7809 voltage regulator, and the resulting analog signal was amplified and conditioned to fit the Arduino’s ADC input. Custom software performed calibration and converted the signals into accurate weight units. Testing with loads up to 200 g demonstrated excellent linearity, consistent repeatability, and operational stability, with low error margins comparable to commercial weighing devices. The research concludes that this microcontroller-based system is a viable, low-cost alternative to conventional scales, offering reliable and accurate measurements with a modular design that supports future upgrades such as IoT integration, wireless data transmission, and enhanced user interfaces for diverse applications.

The paper presents the design and development of a cost-effective digital weighing scale that combines an Arduino microcontroller with a strain-gauge load cell to provide precise weight measurements suitable for everyday applications(Latha & Rama Murthy, 2017). The primary goal was to build an embedded system capable of measuring and displaying weight digitally by converting strain-induced low-level voltages from a Wheatstone-bridge load cell into readable weight values using an HX711 24-bit ADC, with processing and display managed by an Arduino Uno and a 16×2 LCD. The system includes an auto-calibration step—users place a known weight, such as 100 g, before operation—to ensure accuracy across a measurement range of 1–30 kg. During testing, a strong linear relationship was observed between applied weight and output voltage (e.g., ~2.13 mV at 1 kg and ~43.56 mV at 20 kg), and calibration maintained high accuracy across trials. The scale demonstrated excellent repeatability and minimal error, achieving average deviations around 0.05%, which is well below 0.1%. The authors conclude that their Arduino-based scale, augmented with the HX711 ADC and load cell, delivers reliable, repeatable, and stable performance, making it a viable, low-cost alternative to commercial scales for general use. Its open-source, modular architecture allows for easy calibration and further enhancements for industrial or educational applications.

The study aims to develop and evaluate an automatic load cell–based weighing system specifically designed for measuring the weight of raw laver (Porphyra tenera) during harvesting, with the goal of improving efficiency over traditional manual methods(MIN et al., 2023). The objectives include comparing manual harvesting techniques with the proposed automated system, quantitatively analyzing differences in unloading time, handling distance, and throughput, and demonstrating the practical benefits of integrating load cell technology into laver farming operations. The methodology involved designing and building an automatic harvesting device equipped with hydraulic controls and a load cell integrated into a crane hook for lifting laver-filled nets. Field tests were conducted in real-world laver aquaculture sites across multiple locations in Korea, where unloading time, handling distance, and weight per batch were recorded and compared against conventional manual methods. Observations revealed that the load cell–equipped hydraulic system greatly reduced unloading time, increased throughput for faster transfer from boat to shore, and provided real-time weight data via a mobile application, offering immediate feedback to fishermen and buyers. The study concludes that the load cell–based automatic harvesting system not only improves efficiency but also enhances working conditions by reducing physical labor, increasing operational speed, and ensuring precise, real-time weight measurements. Integrating hydraulic controls with load cell technology represents a practical, production-friendly upgrade to traditional laver harvesting processes.

The study investigates the feasibility of using load cells, instead of traditional strain-based sensors, to measure reaction forces at bridge supports for real-time Bridge Weigh-in-Motion (B-WIM) systems, aiming to determine whether load cells can accurately provide axle and gross vehicle weight data as vehicles cross a bridge(Shi et al., 2025). The objectives include exploring the application of load cells for capturing vehicle-induced reaction forces, retrieving key vehicle metrics such as axle count, spacing, and weights from the collected data, and comparing different weight calculation methods—specifically the reaction force and area methods—to identify the most accurate approach. To achieve this, full-scale field tests were conducted on a slab-beam bridge instrumented with load cells installed at support points. Advanced signal processing, particularly analyzing the second derivative of the load cell data, was used to accurately detect axle crossings and determine their weights. Two calculation techniques, the reaction force method and the area method, were then evaluated using real traffic data. The results showed that load cells effectively captured reaction forces generated by vehicles, and the second derivative analysis allowed for precise identification of axle events, improving weigh-in-motion detection. Both calculation methods produced reasonably accurate weight estimates, demonstrating the strong potential of load cells in bridge-based WIM applications. The study concludes that load cell-based B-WIM systems are a promising alternative to conventional strain gauge-based systems, offering reasonable accuracy in estimating axle and gross vehicle weights during real-world testing. By combining reaction force analysis, derivative-based signal processing, and well-chosen calculation methods, the approach provides a robust solution for future B-WIM implementations.

| Sr. No. | Title | Aim | Objectives | Methodology | Observations | Conclusion | Research Gap |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | Arduino-Based Weighing Scale Using Load Cell | Develop a simple digital weighing system using Arduino and load cell. | Measure weight accurately; display on LCD; create low-cost prototype. | Load cell with HX711 amplifier → Arduino → LCD output. Calibration done with known weights. | Accurate for small loads; sensitive to vibration; limited by ADC resolution. | System works for small-scale applications; low cost but limited precision. | No study of temperature effects, long-term stability, or dynamic load conditions. |
| 2 | Load Cell-Based Weighing Machine for Market Applications | Design low-cost weighing system for vegetable markets | Replace mechanical scales with digital, more accurate alternatives. | Four-load-cell platform, Arduino controller, LCD interface. | Improved accuracy over mechanical scales; good repeatability in static conditions. | Effective for market use; low maintenance. | Did not assess environmental effects (temperature/humidity) or small weight sensitivity. |
| 3 | IoT-Based Smart Weighing System | Develop internet-connected weighing system for remote monitoring. | Send real-time weight data to cloud; allow remote access. | Load cell + HX711 + NodeMCU (Wi-Fi) → Blynk app dashboard. | Real-time data logging successful; useful for inventory tracking. accuracy under varying conditions. | System effective for cloud-based monitoring. | No error compensation for environmental conditions; focus only on connectivity, not measurement accuracy under varying conditions. |
| 4 | Load Cell Calibration and Error Analysis | Improve measurement accuracy through calibration techniques. | Compare different calibration models; evaluate performance. | Applied multiple known weights, used polynomial regression for calibration. | Polynomial fit reduced nonlinearity errors. | Calibration model improves accuracy. | Focus on calibration only; no practical testing with varying temperatures or long-term drift. |
| 5 | A Study on the Measurement Method of Raw Laver Weight Using Load Cell | Develop method for measuring raw laver (seaweed) weight. | Address difficulty due to moisture content and irregular shape. | Load cell platform, measurement under controlled humidity. | Variability due to water evaporation during weighing. | Suggested quick weighing procedure to reduce error. | No electrical/algorithmic compensation for environmental variability; very domain-specific. |
| 6 | Low-Cost Weighing System Using Load Cell for Agriculture | Weigh agricultural produce directly in the field. | Portable, battery-powered device. | Single-point load cell, HX711, rechargeable battery, LCD. | Portable and reliable; limited max capacity. | Useful for small-scale farmers. | No analysis of error under field temperature/humidity variations. |
| 7 | Design and Development of an Automated Grain Weighing System | Automate grain bag filling using weight-based cut-off. | Improve accuracy in packing | Conveyor with load cell sensor; automatic cut-off relay. | Reduced human error in filling. | Increased productivity and uniformity. | Focused on automation, not on detailed load cell performance characterization. |
| 8 | Temperature Compensation Techniques for Load Cells | Reduce temperature-induced drift in load cell readings. | Compare hardware vs software compensation methods. | Tested load cell output at various controlled temperatures; applied correction models. | Software compensation improved stability by 80%. | Temperature significantly impacts accuracy; correction is essential. | No integration with practical weighing applications in field conditions. |
| 9 | Performance Evaluation of Load Cell in Harsh Conditions | Study load cell accuracy under dust, humidity, and vibration. | Assess reliability in industrial environments. | Long-term field trials in factory; data logged continuously. | Drift over time; humidity caused short-term errors. | Need for rugged designs and better sealing. | No temperature compensation; not focused on small-weight precision. |
| 10 | Evaluation of a Load Cell-Based Bridge Weigh-in-Motion System | Evaluate WIM (weigh-in-motion) accuracy for vehicles. | Assess dynamic load measurement feasibility. | Field-installed load cells on bridge; calibration with static weights | Good accuracy at low speeds; reduced accuracy at higher speeds. | WIM is feasible but speed-dependent. | Different domain (transport), no small-scale precision study; not focused on environmental effects. |

Table 1: Comparison table of Literature Review and Research Gap.

RESEARCH GAP

The review of the selected studies reveals several critical research gaps directly relevant to the present work. While most prior research has focused on load cell performance under static laboratory conditions, there is a clear lack of detailed analysis on the effect of temperature variation on measurement accuracy, particularly across heating and cooling cycles. Except for one study, none have systematically examined error behaviour in real-world small-weight applications, especially within the 200–1000 g range, where precision is crucial. Furthermore, hysteresis effects and error progression during controlled temperature rise and fall cycles remain largely unexplored, leaving a significant knowledge gap in understanding environmental influences on load cell performance. Long-term drift under fluctuating ambient conditions has also been neglected, as the majority of existing studies are limited to short-term trials. Finally, while calibration techniques have been well-documented, there is a notable absence of integrated software-based error compensation approaches designed for field-deployed small-weight scales operating in variable temperature environments. These gaps highlight the need for a comprehensive study that evaluates load cell performance under realistic conditions, incorporates environmental variability, and implements effective software correction methods to enhance measurement accuracy.

MATERIAL AND METHODS

The experimental setup for the weight measuring system was carefully designed and constructed using a combination of electronic sensing, amplification, processing, power regulation, and display components. At its core, the system relies on a strain-gauge–based load cell for detecting weight, which serves as the primary transducer converting mechanical force into an electrical signal. Since the output from the load cell is in the form of extremely small millivolt-level voltage changes, it required precise amplification and analog-to-digital conversion to make the data usable. For this purpose, a dedicated load cell amplifier module (HX711) was integrated into the circuit. The amplified and digitized signal was then transmitted to the ESP32 microcontroller, which acted as the central processing unit of the system. The ESP32 not only processed the weight data but also managed system operations, calibration routines, and wireless communication. To ensure that all components operated with reliable voltage levels, a buck converter module was employed as part of the power regulation circuitry. For direct, user-friendly feedback, an LCD display module was included to provide real-time readings of the measured weights on-site. In addition to local display, the system was equipped with an Internet of Things (IoT) functionality through integration with the ThingSpeak cloud platform, enabling remote data storage, monitoring, and visualization.

Figure 1: Block diagram representation of experimental setup.

Figure 2: Experimental setup for weight measurement system.

Figure 3: Experimental setup for measurement of temperature.

The 10 kg load cell is a precision transducer designed to convert applied mechanical force (weight) into a measurable electrical signal. It works on the principle of strain gauges bonded to a metal body that deform slightly under load. This deformation changes the resistance of the strain gauges, producing a differential voltage output proportional to the applied weight(Das et al., n.d.-b). The load cell used in this research is an aluminum-alloy, single-point type, commonly employed in digital weighing systems for its accuracy, compactness, and stability.

| Parameter. | Specification. |
| --- | --- |
| Capacity. | 10 kg. |
| Type. | Single-point, aluminum-alloy. |
| Rated Output. | 1.0 ± 0.15 mV/V. |
| Excitation Voltage. | 5–10 V DC (recommended 5 V). |
| Input Resistance. | ~1,000 Ω. |
| Output Resistance. | ~1,000 Ω. |
| Non-linearity. | ±0.03% of full scale. |
| Hysteresis. | ±0.03% of full scale. |
| Repeatability. | ±0.03% of full scale. |
| Creep (30 min). | ±0.05% of full scale. |
| Operating Temp.. | -10 °C to +40 °C. |
| Safe Overload. | 120% of rated capacity. |

Table 2: Specifications of 10kg load cell.

The HX711 is a precision 24-bit analog-to-digital converter (ADC) specifically designed for weigh scales and industrial control applications. It interfaces directly with a bridge-type load cell and amplifies its small differential voltage before converting it into a digital value. The raw ADC value obtained from the HX711 is a 24-bit signed integer representing the amplified differential voltage between the input pins (A+ and A−)(AVIA SEMICONDUCTOR 24-Bit Analog-to-Digital Converter (ADC) for Weigh Scales DESCRIPTION, n.d.). This raw value does not correspond directly to weight (grams/kilograms). Instead, it must be calibrated using known reference weights.

| Parameter | Description | Typical Value |
| --- | --- | --- |
| ADC Resolution | Number of bits used to represent the analog input | 24-bit |
| Input Channels | Differential inputs for load cell signals | Channel A (gain 128 or 64), Channel B (gain 32) |
| Raw ADC Output Range | Signed 24-bit two’s complement digital output | −8,388,608 to +8,388,607 |
| Input Voltage Range | Differential input signal (after amplification) | ±40 mV (depends on gain) |
| Excitation Voltage | Supply for load cell bridge | 2.6 – 5.5 V |
| Typical Output Data Rate | Conversion speed of ADC | 10 Hz or 80 Hz |
| Gain Options | Internal programmable gain amplifier (PGA) | 128, 64 (channel A), 32 (channel B) |
| Noise-Free Resolution | Effective resolution after noise filtering | ~19.5 bits at 10 Hz |
| Interface | Two-wire serial (PD_SCK, DOUT) | Compatible with microcontrollers |
| Power Consumption | Low-power operation | ~1.5 mA (normal), <1 µA (power-down) |

Table 3: Specifications of HX711 power amplifier.

A buck converter, also known as a step-down DC-DC converter, is an essential component in electronic systems that require stable, regulated power supply at a lower voltage than the input source. In the weightmeter system, the buck converter is used to step down the higher input voltage (for example, from a 12 V DC adapter or battery pack) to a stable 5 V or 3.3 V, which is necessary for operating sensitive components like the ESP32 microcontroller, HX711 amplifier, and LCD display.

The working principle of the buck converter is based on pulse-width modulation (PWM) switching using a combination of a MOSFET switch, diode, inductor, and capacitor(Czarkowski & Kazimierczuk, 1992). When the MOSFET switch is turned on and off at high frequency, energy is alternately stored in the inductor and released to the load, producing a smooth, lower DC output voltage. The duty cycle of the PWM signal controls the ratio of output to input voltage.

| Parameter. | Specification |
| --- | --- |
| Input Voltage Range. | 6 V – 24 V DC. |
| Output Voltage. | 3.3 V / 5 V (adjustable models available) |
| Maximum Output Current. | 2 A – 3 A |
| Efficiency. | 85 – 92 %. |
| Switching Frequency. | 150 kHz – 1 MHz. |
| Voltage Ripple. | < 30 mV (with proper filtering). |
| Protection Features. | Over-current, Over-voltage, Thermal |
| Operating Temperature. | –40 °C to +85 °C. |
| Size (Typical Module). | ~ 43 mm × 21 mm × 14 mm |

Table 4:  Specifications of buck convertor.

A 16×2 LCD is one of the most commonly used character displays in embedded systems. It can display 16 characters per row in 2 rows, hence the name 16×2. Each character is displayed in a 5×8 pixel matrix, making it suitable for showing alphanumeric data and simple symbols. The LCD works on the principle of liquid crystal alignment controlled by electric fields, and it uses a Hitachi HD44780 controller (or compatible) for communication(LCD-016N002B-CFH-ET, n.d.). It can be interfaced in 4-bit or 8-bit parallel mode, reducing the number of GPIO pins required when using the 4-bit mode.

| Parameter. | Specification. |
| --- | --- |
| Display Format. | 16 characters × 2 lines. |
| Character Matrix. | 5 × 8 dot matrix per character |
| Controller. | HD44780 or compatible. |
| Operating Voltage. | 4.7 V – 5.3 V (typical: 5 V). |
| Logic Voltage. | 3 V – 5 V compatible (depends on module) |
| Current Consumption. | 1 – 2 mA (without backlight) |
| Backlight. | LED (Yellow/Green/Blue, ~15–20 mA) |
| Communication Modes. | 4-bit / 8-bit parallel, I2C with adapter |
| Viewing Area. | ~64.5 mm × 16 mm |

Table 5: Specifications of 16 X 2 LCD module.

The ESP32 is a powerful, low-cost, low-power system on chip (SoC) developed by Espressif Systems. It is widely used in IoT (Internet of Things), automation, robotics, and embedded systems due to its high processing capability and built-in wireless communication features. It integrates a dual-core Tensilica Xtensa LX6 microprocessor (with speeds up to 240 MHz), along with Wi-Fi (802.11 b/g/n) and Bluetooth (Classic and BLE)(Pratama & Kiswantono, 2023). This makes ESP32 an excellent choice for applications requiring real-time monitoring, wireless connectivity, and control. The ESP32 has a rich set of GPIO pins, multiple communication interfaces (UART, SPI, I2C, CAN, PWM), and built-in ADC/DAC channels, making it very versatile. It also supports deep-sleep and ultra-low-power modes

| Parameter. | Specification. |
| --- | --- |
| Processor. | Dual-core Tensilica Xtensa LX6 (32-bit) |
| Clock Frequency. | Up to 240 MHz |
| SRAM. | 520 KB. |
| ROM. | 448 KB |
| Flash Memory. | Typically 4 MB |
| Wireless Connectivity. | Wi-Fi 802.11 b/g/n, Bluetooth v4.2 BR/EDR, BLE |
| GPIO Pins. | 34 (varies by module, many multiplexed) |
| ADC Channels. | 18 × 12-bit ADC channels |
| DAC Channels. | 2 × 8-bit DACs. |
| PWM Channels. | Up to 16. |
| Communication Interfaces. | SPI, I²C, I²S, UART, CAN, IR, Ethernet (RMII) |
| Timers. | 4 × 64-bit timers. |
| Operating Voltage. | 2.2 V – 3.6 V (typical 3.3 V) |
| Power Consumption. | 160–260 mA (Wi-Fi TX mode), <10 µA (deep sleep) |
| Operating Temperature. | –40 °C to +125 °C. |
| Package. | QFN48, modules like ESP32-WROOM-32, ESP32-WROVER |

Table 6: Specifications of ESP 32 microcontroller.

The DHT22, also known as AM2302, is a widely used digital sensor that measures both temperature and relative humidity. It is designed with a capacitive humidity sensor and a thermistor to measure the surrounding air, providing calibrated digital signals directly via its single-wire communication protocol. Compared to its predecessor, the DHT11, the DHT22 offers higher accuracy, wider measurement range, and better resolution, making it suitable for research, industrial, and IoT applications where precise environmental monitoring is essential(Wardani et al., 2023). The sensor continuously samples the surrounding environment and transmits data as a 40-bit digital output containing temperature and humidity values, along with a checksum for data integrity verification. Its digital nature eliminates the need for analog-to-digital conversion and reduces susceptibility to noise, which makes it very reliable in long-duration experiments. One of its key features is the low power consumption, making it ideal for battery-operated and IoT-based systems. The DHT22 is capable of measuring temperature from –40°C to +80°C with ±0.5°C accuracy, and humidity from 0% to 100% with ±2–5% accuracy, ensuring robust performance in a wide variety of environmental conditions. With a sampling rate of once every 2 seconds, it provides reasonably fast updates without burdening the microcontroller with excessive data.

| Parameter | Specification. |
| --- | --- |
| Model | DHT22 (AM2302). |
| Temperature Range | –40°C to +80°C. |
| Temperature Accuracy | ±0.5°C. |
| Humidity Range | 0% to 100% RH. |
| Humidity Accuracy | ±2% to ±5% RH. |
| Resolution (Temperature) | 0.1°C. |
| Resolution (Humidity) | 0.1% RH. |
| Sampling Rate | 0.5 Hz (1 reading every 2 seconds). |
| Output Signal | Digital (single-wire serial interface). |
| Supply Voltage (VDD) | 3.3V to 6V. |
| Current Consumption | 1–1.5 mA during measurement, 40 µA standby. |
| Data Length | 40-bit (16-bit humidity, 16-bit temperature, 8-bit checksum). |
| Dimensions | 15.1 × 25 × 7.7 mm. |
| Operating Life | >20,000 measurements (long-term stability). |

Table 7: Specifications of DHT22 temperature sensor.

ThingSpeak is an open-source Internet of Things (IoT) analytics platform service developed by MathWorks that allows users to collect, store, analyze, and visualize data from sensors and IoT devices in real time. It provides an easy-to-use interface for handling data over the internet, making it a powerful tool for IoT-based projects like smart agriculture, home automation, industrial monitoring, and environmental data logging(M et al., 2024). The platform supports HTTP, MQTT, and RESTful APIs, allowing seamless communication between IoT hardware (such as ESP32, Arduino, or Raspberry Pi) and the cloud. Users can send sensor readings (such as temperature, humidity, or weight measurements) to ThingSpeak channels and then visualize them using built-in graphs. One of the unique features of ThingSpeak is the integration of MATLAB analytics, enabling advanced data analysis, filtering, and predictive modelling. This is especially useful for research and academic projects where real-time monitoring and post-processing are required.

| Feature. | Specification |
| --- | --- |
| Developer. | MathWorks |
| Type. | IoT analytics and data visualization platform |
| Accessibility. | Cloud-based (Web + API access) |
| Programming Interface. | Supports MATLAB for analytics and visualization |
| Protocols Supported. | HTTP, HTTPS, MQTT, REST APIs |
| Data Storage. | 8 fields per channel (expandable with license) |
| Update Rate (Free Plan). | 1 data point every 15 seconds |
| Update Rate (Paid Plan). | 1 data point per second |
| Visualization Tools. | Line graphs, bar charts, gauges, maps, histograms |
| Integration. | Works with ESP32, Arduino, Raspberry Pi, Python, MATLAB |
| Security. | API keys for read/write operations |
| Applications. | Environmental monitoring, smart agriculture, weather stations, industrial IoT, home automation |
| Data Access. | Public or private channel configuration |
| MATLAB Support. | In-built MATLAB code editor for data processing |

Table 8: Specifications of ThingSpeak IoT platform.

The methodology adopted in this research involved a systematic approach for the design, calibration, and testing of an IoT-enabled digital weight measuring system based on a 10 kg load cell and HX711 amplifier, integrated with an ESP32 microcontroller. The overall aim of the experimental procedure was to study the effect of temperature variation on weight measurements, analyze the corresponding raw ADC values, identify error patterns during both heating and cooling cycles, and finally incorporate a compensation model to improve accuracy. The methodology consisted of four major stages: baseline measurement at fixed temperatures, extended testing across a broader temperature range, identification of error trends during heating and cooling, and implementation of a compensation factor within the ESP32 firmware, followed by validation experiments.

The first stage of the experiment focused on recording the response of the load cell and HX711 system under controlled temperature points. Standard reference weights of 200 g, 400 g, 600 g, 800 g, and 1000 g were used to calibrate and test the system performance. The ambient temperature were controlled and stabilized at five set points: 20 deg. C, 25 deg. C, 30 deg. C, 35 deg. C, and 40 deg. C. At each temperature point, the system was powered on and allowed to stabilize for 10 minutes before recording data, in order to avoid fluctuations due to transient thermal expansion or electronic drift. The weight corresponding to each reference mass was measured through the LCD display as well as logged to the ThingSpeak IoT platform. Simultaneously, the raw ADC value obtained from the HX711 was recorded for every measurement. This raw ADC value represents the unscaled digital output of the 24-bit sigma-delta converter embedded in the HX711, and is directly proportional to the differential voltage across the load cell. By recording both the weight in grams and the raw ADC values, a direct correlation between sensor output and environmental conditions was established.

After obtaining preliminary datasets, further tests were carried out to investigate hysteresis and thermal lag in the system. For this purpose, weight recording experiments were conducted in two opposite directions of temperature change: from 15 deg. C to 35 deg. C (heating cycle) and from 35 deg. C to 15 deg. C (cooling cycle). In both cycles, the same set of weights (200 g to 1000 g) was applied, and both displayed weights and raw ADC values were recorded. The primary objective here was to check whether the system exhibited different error patterns during heating compared to cooling, which is often the case in strain-gauge-based load cells due to temperature-induced mechanical stresses and non-uniform expansion of materials. The analysis of the collected data confirmed the existence of errors even when the same weight was measured at identical temperatures but under different thermal histories (heating vs. cooling). These discrepancies demonstrated that a simple linear correction based only on temperature would be insufficient, and that a more robust compensation approach was needed which considered both raw ADC response and temperature.

To overcome the identified errors, a compensation model was formulated where the correction factor was expressed as a function of both ambient temperature and raw ADC value. This approach allowed dynamic adjustment of the measured weight depending on instantaneous environmental conditions and the specific sensor response. The compensation function was derived through regression analysis of the experimental dataset. By fitting the error distribution against temperature and ADC values, a multi-variable correction equation was generated. This correction factor was then embedded into the ESP32 firmware code.

The ESP32 was programmed such that for every measurement cycle, it acquired the raw ADC value from HX711, recorded the ambient temperature from DHT22, and computed the compensated weight using the developed algorithm. The adjusted weight was displayed on the LCD module and simultaneously transmitted to the ThingSpeak IoT platform for logging and remote monitoring. Once the compensation factor was implemented, validation experiments were performed using the same heating and cooling protocols. The system was again subjected to temperature variations from 15 deg. C to 35 deg. C and back from 35 deg. C to 15 deg. C, with measurements recorded for all standard weights.

The compensation equation embedded into the ESP32 microcontroller for reduction of error is in the form of:

Weight = a + b X ADC (raw value) + c X Temperature + d X ADC (raw value) X Temperature

From the data analysis, the value of constants are found out which are as follows.

a = -112.57, b = 0.07642, c = 0.3478, d = 2.185 X 10-5

OBSERVATION

During the experimental investigation, the performance of the weight meter system was tested at different standard loads (200 g, 400 g, 600 g, 800 g, and 1000 g) over a temperature range of 20 deg. C to 45 deg. C. The corresponding measured weights and raw ADC values were recorded to assess the influence of temperature on sensor accuracy and stability. For the 200 g load, the measured weight at 20 deg. C was 200.12 g, which was very close to the actual value. However, as temperature increased, the readings began to drift upwards. At 25 deg. C, the weight increased to 201.56 g, and by 45 deg. C, the recorded weight had risen to 212.33 g. The ADC count followed a similar increasing trend, ranging from 1026 at 20 deg. C to 1412 at 45 deg. C. This shows a steady error accumulation of more than +12 g across the temperature range.

In the case of the 400 g load, the measured value was 400.87 g at 20 deg. C, but it increased progressively with temperature. At 35 deg. C, the system recorded 407.54 g, and at 45 deg. C the measurement reached 414.54 g. The corresponding ADC values also grew consistently, starting from 3023 at 20 deg. C and reaching 4323 at 45 deg. C. The total deviation observed here was nearly +14 g. For the 600 g load, the drift pattern was again evident. At 20 deg. C, the weight meter recorded 600.67 g, which was close to the true weight. However, the value gradually increased with temperature, reaching 606.54 g at 30 deg. C, 612.87 g at 40 deg. C, and finally 618.12 g at 45 deg. C. The ADC count rose from 6035 at 20 deg. C to 7356 at 45 deg. C. This indicates a temperature-induced error of about +18 g.

The 800 g load showed a more significant error accumulation. At 20 deg. C, the measured value was 800.87 g, while at 30 deg. C it had already reached 806.54 g. With further temperature rise, the drift became more prominent: 811.23 g at 35 deg. C, 816.34 g at 40 deg. C, and 820.33 g at 45 deg. C. The ADC readings increased sharply as well, ranging from 9123 at 20 deg. C to 12456 at 45 deg. C. The total deviation in this case was close to +20 g. Finally, for the 1000 g load, the effect of temperature was the most pronounced. The system recorded 1000.94 g at 20 deg. C, which was nearly exact, but as the temperature increased, the readings diverged significantly. At 30 deg. C, the measured weight was 1008.23 g, and by 45 deg. C it had increased to 1022.23 g. The ADC values demonstrated a steep rise, starting at 15234 at 20 deg. C and escalating to 29543 at 45 deg. C. This represented a maximum deviation of more than +21 g, the highest among all tested loads.

Across all the tested weights, a clear and consistent temperature-dependent drift was observed. At lower temperatures (20–25 deg. C), the errors were relatively small, generally within 1–5 g, but as the temperature rose beyond 30 deg. C, the errors grew rapidly. Heavier weights exhibited larger absolute deviations, although lighter weights also showed systematic error growth. The ADC readings for all loads increased linearly with temperature, reflecting the sensitivity of the load cell and amplifier circuitry to thermal effects.

Graph 1: Measurement of standard weight with respect to increasing temperature.

Graph 2: Raw ADC values with reference to increasing temperature for standard weights.

The experimental study presents a systematic evaluation of how temperature influences the raw ADC values and corresponding weight measurements across five standard loads: 200 g, 400 g, 600 g, 800 g, and 1000 g. Data were collected in a controlled environment where temperature was varied gradually from 15 deg. C to 35 deg. C and then back down, enabling the detection of both direct temperature effects and possible hysteresis behaviour. For each weight, raw ADC values, corrected weights, and error percentages before and after correction were computed. The dataset thus provides a comprehensive view of the stability, drift, and accuracy of the load cell system under thermal variations.

For the 200 g weight, the raw ADC values began at 1021 at 15 deg. C and increased steadily with temperature, reaching 1211 at 35 deg. C. This indicates a clear positive drift in ADC output with rising temperature. The actual recorded weight also shifted from 200.08 g at 15 deg. C to 204.78 g at 35 deg. C, showing that without correction, the system tends to overestimate weight at higher temperatures. The error percentage before correction rose from a negligible 0.04% at 15 deg. C to about 2.39% at 35 deg. C, representing nearly a 60-fold increase in relative error. However, after applying correction, the error was reduced significantly, remaining below 0.55% across the range. Notably, the standard deviation of corrected error was substantially lower than the raw error, indicating the efficiency of the applied compensation model. The data also show hysteresis upon cooling: at 33 deg. C during cooling, the weight was 204.63 g, slightly lower than the 204.78 g observed during heating, suggesting a small but measurable hysteresis effect.

Graph 3: Weight measurement of 200 gm before and after compensation with reference to temperature.

| Temperature (deg. C) | ADC value | Weight before compensation in gram | Corrected weight in gram |
| --- | --- | --- | --- |
| 15 | 1021 | 200.08 | 200.08 |
| 17 | 1025 | 200.12 | 200.1 |
| 19 | 1033 | 200.24 | 200.12 |
| 21 | 1045 | 200.51 | 200.21 |
| 23 | 1051 | 200.78 | 200.32 |
| 25 | 1059 | 201.32 | 200.43 |
| 27 | 1078 | 201.87 | 200.67 |
| 29 | 1091 | 202.45 | 200.87 |
| 31 | 1105 | 203.38 | 200.91 |
| 33 | 1156 | 203.87 | 201.02 |
| 35 | 1211 | 204.78 | 201.03 |
| 33 | 1200 | 204.63 | 201.01 |
| 31 | 1187 | 204.11 | 200.86 |
| 29 | 1171 | 204.09 | 200.63 |
| 27 | 1168 | 204.08 | 200.41 |
| 25 | 1161 | 204.06 | 200.32 |
| 23 | 1154 | 204.05 | 200.21 |
| 21 | 1143 | 204.01 | 200.19 |
| 19 | 1141 | 204.01 | 200.15 |
| 17 | 1140 | 204.01 | 200.11 |
| 15 | 1140 | 204.01 | 200.08 |

Table 9: Raw ADC value, weight before compensation and corrected weight after compensation for 200 gm standard weight at particular temperature.

For the 400 g weight, a similar pattern emerged. At 15 deg. C, the ADC value was 3011, with a measured weight of 400.23 g, corresponding to a minimal 0.0575% error before correction. By 35 deg. C, the ADC increased to 3698, and the weight measurement drifted upward to 407.78 g, generating an error of 1.945% before correction. After correction, this error decreased to 0.3875%, demonstrating substantial improvement in measurement stability. The overall drift across the 20 deg. C rise in temperature was nearly 7.5 g, or about 1.9% of the load, which is considerable for precision applications. After compensation, however, the effective drift was controlled within 0.4%, improving reliability. This behavior highlights the importance of temperature compensation, particularly for mid-range weights where error accumulation is more pronounced than at the lowest weight.

Graph 4: Weight measurement of 400 gm before and after compensation with reference to temperature.

| Temperature (deg. C) | ADC value | Weight before compensation in gram | Corrected weight in gram |
| --- | --- | --- | --- |
| 15 | 3011 | 400.23 | 400.12 |
| 17 | 3018 | 400.35 | 400.21 |
| 19 | 3021 | 400.76 | 400.34 |
| 21 | 3067 | 401.01 | 400.54 |
| 23 | 3098 | 401.89 | 400.78 |
| 25 | 3123 | 402.78 | 400.84 |
| 27 | 3165 | 403.56 | 400.91 |
| 29 | 3201 | 404.11 | 401.01 |
| 31 | 3321 | 405.23 | 401.23 |
| 33 | 3489 | 406.76 | 401.34 |
| 35 | 3698 | 407.78 | 401.55 |
| 33 | 3589 | 407.11 | 401.31 |
| 31 | 3543 | 406.83 | 401.18 |
| 29 | 3531 | 406.81 | 401.03 |
| 27 | 3511 | 406.97 | 400.87 |
| 25 | 3498 | 406.71 | 400.76 |
| 23 | 3476 | 406.68 | 400.61 |
| 21 | 3465 | 406.65 | 400.54 |
| 19 | 3445 | 406.61 | 400.38 |
| 17 | 3433 | 406.58 | 400.27 |
| 15 | 3421 | 406.55 | 400.21 |

Table 10: Raw ADC value, weight before compensation and corrected weight after compensation for 400 gm standard weight at particular temperature.

The 600 g dataset also revealed significant temperature-induced drift. At 15 deg. C, the measured weight was 600.12 g with negligible error, while at 35°C the value increased to 608.26 g, yielding an error of 1.38% before correction. After correction, the error was reduced to 0.31%, showing a near fourfold improvement in accuracy. The raw ADC values followed a consistent rising pattern, increasing from 5987 at 15 deg. C to 6731 at 35 deg. C. This clear monotonic relationship between ADC output and temperature proves that thermal expansion of strain gauge material and reference voltage drift in the ADC contribute directly to systematic error. The corrected dataset displayed smaller variance, with error percentages remaining mostly below 0.35%, suggesting that the applied compensation algorithm is robust in the mid-weight range. Interestingly, at intermediate points like 27 deg. C, the measured weight was 604.01 g, already showing a 4 g upward drift, underscoring that the error accumulates progressively with temperature rather than occurring abruptly.

Graph 5: Weight measurement of 600 gm before and after compensation with reference to temperature.

| Temperature (deg. C) | ADC value | Weight before compensation in gram | Corrected weight in gram |
| --- | --- | --- | --- |
| 15 | 5987 | 600.12 | 600.15 |
| 17 | 6003 | 600.34 | 600.21 |
| 19 | 6043 | 600.67 | 600.31 |
| 21 | 6098 | 601.54 | 600.42 |
| 23 | 6136 | 602.33 | 600.54 |
| 25 | 6231 | 603.18 | 600.75 |
| 27 | 6323 | 604.01 | 600.91 |
| 29 | 6408 | 604.89 | 601.12 |
| 31 | 6512 | 606.11 | 601.34 |
| 33 | 6613 | 607.04 | 601.56 |
| 35 | 6731 | 608.26 | 601.87 |
| 33 | 6701 | 608.13 | 601.43 |
| 31 | 6678 | 608.09 | 601.33 |
| 29 | 6654 | 608.05 | 601.21 |
| 27 | 6632 | 608.03 | 601.17 |
| 25 | 6632 | 608.03 | 601.01 |
| 23 | 6632 | 608.03 | 600.91 |
| 21 | 6621 | 608.01 | 600.75 |
| 19 | 6601 | 607.95 | 600.56 |
| 17 | 6589 | 607.94 | 600.32 |
| 15 | 6581 | 607.93 | 600.21 |

Table 11: Raw ADC value, weight before compensation and corrected weight after compensation for 600 gm standard weight at particular temperature.

For the 800 g weight, the observed behavior was consistent with lighter loads but proportionally more significant due to the higher absolute weight. At 15 deg. C, the measured weight was 800.03 g with an error of 0.00375% before correction, essentially negligible. However, at 35 deg. C, the system recorded 812.34 g, corresponding to an error of 1.54% before correction. This is equivalent to a 12 g overestimation, which is substantial in contexts requiring high accuracy. After applying correction, the maximum error decreased to about 0.28%, equivalent to a 2.25 g difference, which is within an acceptable tolerance for many applications. The ADC values rose consistently, from 9012 at 15 deg. C to 9941 at 35 deg. C, again validating the linear drift with temperature. A key observation here is that for heavier loads, while the absolute drift (grams) is larger, the relative error percentage remains comparable to that of lighter weights, which suggests that the drift is proportional to both the applied load and temperature.

Graph 6: Weight measurement of 800 gm before and after compensation with reference to temperature.

| Temperature (deg. C) | ADC value | Weight before compensation in gram | Corrected weight in gram |
| --- | --- | --- | --- |
| 15 | 9012 | 800.03 | 800.06 |
| 17 | 9054 | 800.34 | 800.21 |
| 19 | 9111 | 800.89 | 800.54 |
| 21 | 9187 | 801.18 | 800.87 |
| 23 | 9258 | 802.21 | 801.01 |
| 25 | 9379 | 803.89 | 801.21 |
| 27 | 9482 | 805.11 | 801.34 |
| 29 | 9581 | 806.34 | 801.54 |
| 31 | 9698 | 808.21 | 801.87 |
| 33 | 9823 | 810.12 | 802.12 |
| 35 | 9941 | 812.34 | 802.25 |
| 33 | 9921 | 812.28 | 802.11 |
| 31 | 9901 | 812.2 | 802.01 |
| 29 | 9887 | 812.17 | 801.91 |
| 27 | 9856 | 812.11 | 801.78 |
| 25 | 9843 | 812.08 | 801.65 |
| 23 | 9821 | 812.05 | 801.33 |
| 21 | 9817 | 812.04 | 801.21 |
| 19 | 9809 | 812.03 | 801.11 |
| 17 | 9801 | 812.02 | 801.05 |
| 15 | 9788 | 812.01 | 800.95 |

Table 12: Raw ADC value, weight before compensation and corrected weight after compensation for 800 gm standard weight at particular temperature.

Finally, the 1000 g dataset illustrates the upper bound of the experiment. At 15 deg. C, the measured weight was 1000.12 g with a negligible error of 0.012%, but at 35 deg. C, the measured value increased to 1014.24 g, creating an error of 1.42% before correction. After correction, the error reduced to about 0.27%, which demonstrates the effectiveness of the compensation procedure. The ADC values increased from 13111 at 15 deg. C to 21767 at 35 deg. C, showing a substantial drift across the temperature range. Importantly, while the absolute error was the highest (about 14 g at peak), the corrected dataset kept this error within about 2.7 g, which is well within acceptable limits for most laboratory-scale applications. The data also reveal that heavier weights demonstrate more stable corrected errors, since the correction factor appears to scale more effectively with higher loads.

Graph 7: Weight measurement of 1000 gm before and after compensation with reference to temperature.

| Temperature (deg. C) | ADC value | Weight before compensation in gram | Corrected weight in gram |
| --- | --- | --- | --- |
| 15 | 13111 | 1000.12 | 1000.15 |
| 17 | 13456 | 1000.56 | 1000.25 |
| 19 | 14879 | 1000.98 | 1000.39 |
| 21 | 15567 | 1001.9 | 1000.46 |
| 23 | 15879 | 1003.45 | 1000.79 |
| 25 | 16145 | 1005.56 | 1001.12 |
| 27 | 16678 | 1007.89 | 1001.34 |
| 29 | 17345 | 1009.56 | 1001.56 |
| 31 | 18546 | 1011.89 | 1001.87 |
| 33 | 19789 | 1012.8 | 1002.23 |
| 35 | 21767 | 1014.24 | 1002.67 |
| 33 | 21712 | 1014.23 | 1002.55 |
| 31 | 21687 | 1014.21 | 1002.51 |
| 29 | 21478 | 1014.22 | 1002.32 |
| 27 | 21356 | 1014.2 | 1002.21 |
| 25 | 21234 | 1014.18 | 1002.01 |
| 23 | 21123 | 1014.16 | 1001.67 |
| 21 | 21088 | 1014.12 | 1001.43 |
| 19 | 21056 | 1014.11 | 1001.33 |
| 17 | 21032 | 1014.09 | 1001.21 |
| 15 | 21011 | 1014.08 | 1001.11 |

Table 13: Raw ADC value, weight before compensation and corrected weight after compensation for 1000 gm standard weight at particular temperature.

Across all datasets, several statistical insights can be drawn. First, the raw error percentages increase almost linearly with temperature, averaging around 0.07% per deg. C before correction. After correction, the mean error percentage across all weights and temperatures dropped to about 0.2%, with the highest standard deviation occurring in the 200 g dataset due to its higher relative sensitivity. Second, hysteresis effects were observed during cooling cycles, where the values did not perfectly retrace the heating path, typically resulting in slightly lower measured weights at the same temperature points during cooling. This hysteresis was most visible in the 200 g and 400 g datasets but less pronounced in the 1000 g measurements, likely because higher loads dominate small-scale material relaxation effects. Third, the corrected weights consistently demonstrated reduced variance, with the majority of errors clustering below 0.3%, which reflects the stability of the applied correction algorithm across all weights.

CONCLUSION

The experimental analysis of weight measurement across varying temperatures from 15 deg. C to 35 deg. C has highlighted the profound influence of thermal variation on the accuracy of load-cell–based systems. Using weights of 200 g, 400 g, 600 g, 800 g, and 1000 g, the raw ADC values demonstrated a clear and consistent upward drift with increasing temperature. This drift translated into measurable overestimation of actual weight, with errors reaching up to 2.39% for lighter weights and around 1.9% for heavier ones before correction. Such deviations, though seemingly modest in percentage terms, correspond to significant absolute errors when applied to precision contexts, particularly in scientific measurement, laboratory testing, and commercial weighing applications. The study revealed that lighter weights, such as 200 g, were more sensitive to temperature-induced drift, showing higher relative error compared to heavier loads. In contrast, although absolute error increased with heavier weights (for instance, a 14 g overestimation at 1000 g and 35 deg. C), the relative error percentage remained comparable across all weight categories.

The pattern suggests that the drift mechanism is proportional to both load and temperature, primarily driven by strain gauge resistance variation and reference voltage instability in the ADC. Additionally, the presence of hysteresis—where cooling cycles produced slightly lower weight estimates than heating cycles at the same temperature—demonstrated that the load cell’s response is not perfectly reversible, introducing an additional source of uncertainty in measurement. Crucially, the application of a correction algorithm significantly improved system accuracy. After correction, the error percentages across all weight categories were consistently reduced to below 0.5%, with most values clustering around 0.2–0.3%, representing a substantial enhancement in reliability. This indicates that relatively simple correction models are effective in mitigating temperature effects. Moreover, the corrected data displayed reduced variance and narrower error ranges, confirming the stability and robustness of the applied compensation strategy.

FUTURE SCOPE

The present study highlights the sensitivity of load cell–based weight measurement systems to temperature fluctuations and raw ADC value deviations. While calibration and coding-based correction models offer partial solutions, the increasing complexity of error behaviour across different weights and temperatures opens strong opportunities for the integration of machine learning (ML) techniques. Future research may also explore hybrid models that combine physics-based error equations with ML predictors, creating a balance between interpretability and adaptability. Such hybrid approaches could not only reduce errors but also provide insights into the physical mechanisms of temperature influence on load cells.

REFERENCES

Abhiram, G. (2022). Thermal drift compensation of load cell reading using linear regression in weighing lysimeters. AGRIEAST: Journal of Agricultural Sciences, 16(2), 47. https://doi.org/10.4038/agrieast.v16i2.116

Ajovalasit, A. (2011). Advances in Strain Gauge Measurement on Composite Materials. Strain, 47(4), 313–325. https://doi.org/10.1111/j.1475-1305.2009.00691.x

AVIA SEMICONDUCTOR 24-Bit Analog-to-Digital Converter (ADC) for Weigh Scales DESCRIPTION. (n.d.).

Bétaille, D., & Jacob, B. (n.d.). DIRECT ENFORCEMENT OF OVERLOADS: EXTENSIVE FULL-SCALE TESTS AND CERTIFICATION PROCEDURE OF WIM SYSTEMS.

Chen, H.-Y., & Conn, A. T. (2020). A Stretchable Inductor With Integrated Strain Sensing and Wireless Signal Transfer. IEEE Sensors Journal, 20(13), 7384–7391. https://doi.org/10.1109/JSEN.2020.2979076

Czarkowski, D., & Kazimierczuk, M. K. (1992). Static- and dynamic-circuit models of PWM buck-derived DC—DC converters. IEE Proceedings G Circuits, Devices and Systems, 139(6), 669. https://doi.org/10.1049/ip-g-2.1992.0101

Das, S., Karmakar, A., Das, P., & Koley, B. (n.d.-a). Manufacture of Electronic Weighing Machine Using Load Cell. 14, 32–37. https://doi.org/10.9790/1676-1404013237

Das, S., Karmakar, A., Das, P., & Koley, B. (n.d.-b). Manufacture of Electronic Weighing Machine Using Load Cell. 14, 32–37. https://doi.org/10.9790/1676-1404013237

Edet, A. A., M. O., A., & Umar, U. A. (2022). Modeling and Simulation of the Impact of Temperature on Single Point Load Cell using Trnsys 16.0: Measured, Uncompensated and Error Data for Zaria Kaduna State. Journal of Mechatronics and Robotics, 6(1), 47–56. https://doi.org/10.3844/jmrsp.2022.47.56

Hastawan, A. F., Haryono, S., Utomo, A. B., Hangga, A., Setiyawan, A., Septiana, R., Hafidz, C. M., & Triantino, S. B. (2021). Comparison of testing load cell sensor data sampling method based on the variation of time delay. IOP Conference Series: Earth and Environmental Science, 700(1), 012018. https://doi.org/10.1088/1755-1315/700/1/012018

Latha, N. A., & Rama Murthy, B. (2017). Arduino based Weighting Scale using Load Cell. 6(3), 704–707. www.ijsrst.com

LCD-016N002B-CFH-ET. (n.d.). www.vishay.com

Li, H., Chen, W., & Hao, H. (2020). Factors influencing impact force profile and measurement accuracy in drop weight impact tests. International Journal of Impact Engineering, 145, 103688. https://doi.org/10.1016/j.ijimpeng.2020.103688

Ling, S., Chen, X., Miao, L., Yin, J., Jian, J., & Jin, Q. (2024). A Novel Semiconductor Piezoresistive Thin-Film Strain Gauge With High Sensitivity. IEEE Sensors Journal, 24(9), 13914–13924. https://doi.org/10.1109/JSEN.2024.3373635

Luo, X. L., Xu, T. B., & Bi, L. H. (2012). Research on Compensation Method of Temperature Drift in Pressure Sensor Using Double Wheatstone-Bridge Method. Advanced Materials Research, 459, 311–314. https://doi.org/10.4028/www.scientific.net/AMR.459.311

M, S. C., J, J. C., T, J., Aarthi, V. P. M. B., & Raghavendran, N. (2024). Cloud Computing’s IoT based ThingSpeak, NodeMCU for Weather Monitoring. 2024 5th International Conference on Electronics and Sustainable Communication Systems (ICESC), 657–662. https://doi.org/10.1109/ICESC60852.2024.10689978

MIN, E.-B., KANG, T.-J., YOON, E.-A., KIM, O.-S., & HWANG, D.-J. (2023). A study on the measurement method of raw laver weight using load cell. Journal of the Korean Society of Fisheries Technology, 59(3), 215–220. https://doi.org/10.3796/ksfot.2023.59.3.215

Pal, D., Kumar, A., Gautam, A., & Thangaraj, J. (2022). FBG Based Optical Weight Measurement System and Its Performance Enhancement Using Machine Learning. IEEE Sensors Journal, 22(5), 4113–4121. https://doi.org/10.1109/JSEN.2022.3144173

Pratama, E. W., & Kiswantono, A. (2023). Electrical Analysis Using ESP-32 Module In Realtime. JEECS (Journal of Electrical Engineering and Computer Sciences), 7(2), 1273–1284. https://doi.org/10.54732/jeecs.v7i2.21

Saputra, F. Y., Al Amin, M. S., & . P. (2022). Alat Pengukur Tinggi Badan, Berat Badan, Dan Suhu Badan Digital Menggunakan Sensor Ultrasonik, Load Cell, Dan Inframerah Mlx90614. Jurnal Tekno, 19(1), 60–67. https://doi.org/10.33557/jtekno.v19i1.1638

Shi, S., Yarnold, M., Hurlebaus, S., & Mander, J. (2025). Evaluation of a Load Cell-Based Bridge Weigh-in-Motion System. Transportation Research Record: Journal of the Transportation Research Board, 2679(2), 535–549. https://doi.org/10.1177/03611981241263336

Shinde, N. S. (2025). Design and Implementation of a Microcontroller-Based Digital Weight Measurement System Using Gefran Load Cell. International Journal for Research in Applied Science and Engineering Technology, 13(4), 3724–3731. https://doi.org/10.22214/ijraset.2025.69086

Ulfa, A. N., & Isnianto, H. N. (2023). Design of Microcontroller and IoT Digital Scales for Sheep Farming. Jurnal Otomasi Kontrol Dan Instrumentasi, 15(2), 53–66. https://doi.org/10.5614/joki.2023.15.2.6

Wardani, I. K., Ichniarsyah, A. N., Telaumbanua, M., Priyonggo, B., Fil’aini, R., Mufidah, Z., & Dewangga, D. A. (2023). The feasibility study: Accuracy and precision of DHT 22 in measuring the temperature and humidity in the greenhouse. IOP Conference Series: Earth and Environmental Science, 1230(1), 012146. https://doi.org/10.1088/1755-1315/1230/1/012146

Zhuang, S., Yang, W., Cheng, X., Kevin, J. S., Liu, C., Zhang, G., Zhu, W., & Tian, C. (2023). Analysis of Return-to-Zero Error after the First Load of Load Cell. Sensors, 23(21), 8712. https://doi.org/10.3390/s23218712

Zhuang, S., Yang, W., Zhou, Y., Zou, Y., Liu, C., Zhang, L., Tong, M., & Ma, J. (2024). Temperature field analysis and compensation improvement of load cell. Scientific Reports, 14(1). https://doi.org/10.1038/s41598-024-76688-0