Review Article

Effectiveness of Metal Artifact Reduction Algorithms in Dental CBCT: An Overview

Abstract: Dental cone-beam computed tomography (CBCT) has significantly enhanced diagnostic capabilities in maxillofacial imaging, offering high-resolution, three-dimensional views with relatively low radiation exposure. However, the presence of metallic objects such as implants, restorations, and orthodontic appliances often introduces image-degrading artifacts that compromise diagnostic accuracy. Various metal artifact reduction (MAR) algorithms have been developed to counteract these limitations. This review synthesizes current evidence on the diagnostic effectiveness of MAR algorithms in dental CBCT, emphasizing their methodological principles, clinical outcomes, and limitations across different imaging contexts.

Keywords: CBCT, metal artifacts, MAR algorithms, image quality, diagnostic accuracy

# Introduction:

Cone-beam computed tomography (CBCT) has become an indispensable imaging modality in modern dental practice, providing high-resolution three-dimensional representations of maxillofacial structures with relatively low radiation exposure compared to traditional medical computed tomography (CT) (Hartshorne, 2022). Compared with CT, CBCT generally delivers a lower radiation dose when producing images for equivalent diagnostic objectives. Despite these advantages, one of the main limitations of CBCT is its sensitivity to metal-induced artifacts, especially in patients with existing dental restorations or prosthetic materials (Heller et al., 2024). High-density objects such as titanium implants, zirconia-based crowns, amalgam restorations, and metallic endodontic posts often introduce image distortions (Spagnuolo, 2022). These artifacts are primarily the result of beam hardening and x-ray scatter, phenomena that alter the quality and consistency of the x-ray beam as it passes through dense materials. The consequences are often visible as streak artifacts, dark bands, and non-uniformities in the reconstructed image volume, which can obscure anatomical details or mimic pathology.1

Metal artifacts in CBCT images often present as streaks, as well as areas of exaggerated brightness or darkness, particularly around high-density materials like dental implants or metallic restorations. These visual distortions can obscure critical anatomical landmarks, making it difficult for clinicians to accurately assess bone structures or adjacent soft tissues. Additionally, non-uniform shading caused by beam hardening and scatter may conceal fine bone details, such as cortical boundaries or subtle bone loss. Even more concerning, spurious bands and distortions may resemble pathological conditions, such as peri-implant defects or bone dehiscence, potentially leading to false-positive diagnoses. As a result, the diagnostic reliability of CBCT imaging is significantly compromised in the presence of metal, and such artifacts can ultimately increase the risk of misdiagnosis, particularly in cases where precise visualization is critical for treatment planning or surgical intervention.2

To address the diagnostic challenges posed by metal artifacts, many Cone-Beam Computed Tomography (CBCT) systems now incorporate Metal Artifact Reduction (MAR) algorithms aimed at improving image quality in the presence of metallic structures. These algorithms work by reducing the streaking, beam hardening, and scattering effects that distort CBCT images when high-density materials such as dental implants, crowns, or restorations are present.

Early-generation MAR techniques primarily relied on simple interpolation methods or post-processing filters applied to raw projection data. These methods attempted to fill in missing or corrupted data caused by metal, either by averaging adjacent projections or applying smoothing filters to reduce visible distortions. Examples of such commercially available systems include Picasso Master 3D and ProMax 3D, which integrated basic MAR tools into their imaging workflows. While these approaches offered some degree of visual improvement, they were often limited in effectiveness, especially when metal artifacts were severe or when high diagnostic precision was required.

Advancements in image processing have since led to the development of more sophisticated MAR techniques. Iterative reconstruction algorithms, for example, use repeated cycles of forward and backward projection to refine the image and compensate for data inconsistencies introduced by metal. These methods are more computationally intensive but offer significantly improved artifact suppression and image clarity.

More recently, the integration of artificial intelligence (AI)—particularly deep learning techniques—has marked a major step forward in MAR capabilities. AI-based models can be trained on large datasets to recognize and differentiate artifacts from true anatomical structures, allowing for more intelligent and adaptive correction. These systems have demonstrated promising results in reducing artifacts without compromising the visibility of surrounding tissues, offering a new frontier in diagnostic imaging.

Together, these evolving MAR strategies reflect the ongoing effort to enhance CBCT image quality and ensure more accurate diagnoses in patients with metallic dental materials.3

Metal artifact reduction (MAR) techniques for dental cone-beam computed tomography (CBCT) vary widely in approach and complexity. Some methods focus on modifying the sinogram—the raw projection data—using interpolation or path-length weighting to correct inconsistencies caused by metallic objects. Others segment the metal regions and apply corrective reconstructions, or use pixel-domain filters to reduce artifacts in the reconstructed images. Each approach has its strengths and limitations, and the choice of method often depends on the specific clinical scenario and the severity of the metal artifacts.4

In recent years, various post-processing methods have been developed to minimize the impact of metal artifacts in cone-beam computed tomography (CBCT). One such approach is the auto-edge counting method, which applies a Canny edge detection algorithm to objectively quantify the severity of streak artifacts caused by metallic restorations. This method has been shown to reliably assess artifact burden while reducing the limitations of conventional gray-scale–based evaluations. Experimental work using implants and intracanal posts demonstrated that artifact intensity is significantly influenced by the field of view size and by whether the metal artifact reduction (MAR) algorithm is activated. Specifically, smaller fields of view and MAR-enabled reconstructions were associated with reduced artifact formation, whereas changes in object position and tube voltage had minimal effect. These findings highlight the importance of optimizing acquisition parameters in CBCT to enhance image quality and diagnostic accuracy in cases where high-density dental materials are present.5

Advanced approaches now even use deep learning techniques like Convolutional Neural Networks (CNNs) and Generative Adversarial Networks (GANs). These advanced methods can significantly reduce artifacts such as streaks caused by metal objects, improving image quality across various scan conditions. CNNs are trained to identify and remove metal-induced distortions while preserving anatomical details, making them highly effective for clinical use. GANs, on the other hand, generate artifact-free images by learning from both corrupted and clean data. As computational power continues to increase, these deep learning approaches are poised to become standard in routine CBCT imaging, offering faster, more accurate diagnostics.7

Despite these advances, the clinical effectiveness of MAR in dental CBCT remains unclear. Many studies have evaluated MAR performance in vitro, using phantoms or extracted specimens with controlled defects. Some report visually cleaner images when MAR is enabled, but others find little or no gain (or even harm) in diagnostic accuracy. This review summarizes the evidence on MAR effectiveness in dental CBCT, organized by application (endodontic, implant, phantom) and algorithm type, to guide clinicians on when MAR may be beneficial.

Types of Metal Artifacts in CBCT:

Metallic restorations, implants, and endodontic posts often lead to characteristic artifacts in CBCT imaging, which degrade diagnostic quality and limit interpretation. These artifacts can be broadly classified as follows:

1.Streak artifacts:

- Appear as alternating bright and dark linear streaks radiating from metallic objects.

- Result mainly from severe beam hardening, where lower-energy photons are absorbed while higher-energy photons pass through, leading to an imbalance in the reconstructed image.

- Particularly problematic in endodontically treated teeth with posts or in regions with multiple implants, where streaks obscure periapical and periodontal structures.

2.Shading (cupping) artifacts:

- Manifests as non-uniform darkening around metallic objects, giving a "cupping" or “capping” effect.

- Caused by differential attenuation of X-ray photons within the volume, leading to reduced gray values and contrast in surrounding tissues.

- Can make it difficult to distinguish subtle changes in bone density adjacent to implants or teeth.

3.Dark bands:

- Appear as localized, elongated zones of low attenuation or signal voids adjacent to metallic objects.

- Often arise from truncation artifacts (limited detector size cutting off parts of the projection data) or saturation effects in the detector.

- Clinically, these may be misinterpreted as peri-implant bone defects or bone loss, leading to diagnostic errors.

4.Scatter noise:

- Appears as diffuse haziness or fog-like background across the image.

- Produced by scattered X-rays that deviate from their original paths and are still detected, thereby degrading image contrast.

- Scatter reduces visibility of fine trabecular patterns, small lesions or early periapical changes.

5.Distortions:

- Geometric deformations of anatomical structures near metallic objects.

- Can result from nonlinear effects in image reconstruction algorithms when processing corrupted raw data.

- These distortions can compromise linear measurements (e.g., bone height or width), potentially affecting implant planning or surgical navigation.8

MAR Techniques and Approaches:

A range of Metal Artifact Reduction (MAR) algorithms have been developed, each targeting the specific mechanisms underlying artifact formation. These approaches can be broadly categorized into projection-domain corrections, sinogram correction methods, iterative reconstruction techniques and deep learning–based models.

1. Projection-domain correction:

- These methods operate directly on the raw projection data (sinograms) before image reconstruction.

- The key strategy involves identifying regions affected by metal traces in projection images and then interpolating or replacing corrupted data.

- Early implementations were based on relatively simple interpolation techniques, where missing or degraded projection values caused by metal were substituted with surrounding values.

- Commercial CBCT units such as Vatech’s Picasso Master 3D incorporated this form of MAR, applying correction during the reconstruction phase to minimize streaking and banding around implants and restorations.

- The advantage of this approach lies in its computational efficiency and ability to preserve image sharpness. However, it can sometimes lead to over-smoothing and partial loss of anatomical detail if interpolation is excessive.4,1

2. Sinogram-weighting and Direct Sinogram Correction (DSC):

- Direct Sinogram Correction (DSC) methods focus on corrupted regions of the sinogram and replace them with corrected estimates, thereby improving the quality of the reconstructed volume.

- A significant advancement was reported by Hegazy et al. (2023), who combined DSC with metal path-length weighting. In this approach, sinogram data were weighted proportionally to the amount of metal traversed by the X-ray beam, giving higher priority to less-corrupted rays.

- Their study, performed on both phantom and clinical patient scans at 90 kVp, demonstrated that weighted DSC provided superior artifact suppression compared to conventional DSC, particularly in cases with severe metallic burden.

- This modification enhanced diagnostic clarity without excessively blurring fine details, making it a promising refinement for clinical MAR applications.4

3. Iterative reconstruction approaches:

- Iterative algorithms reconstruct images by simulating the underlying physics of X-ray interaction with matter, including beam hardening, scatter, and noise.

- They rely on repeated segmentation–correction–reprojection loops, in which metallic objects are segmented, their effects estimated, and corrected projections are iteratively refined.

- Hsieh et al. (2000) proposed an iterative approach specifically for beam hardening correction in cone-beam CT, which highlighted how iterative re-estimation of the X-ray attenuation process could reduce streak artifacts.

- While computationally more intensive than direct interpolation or DSC, iterative methods have the advantage of producing more physically consistent reconstructions with fewer distortions. However, their performance is highly dependent on accurate segmentation of metallic regions.9

4. Deep learning–based MAR:

- More recently, Artificial Intelligence (AI), particularly deep learning, has been applied to MAR.

- Convolutional Neural Networks (CNNs) and Generative Adversarial Networks (GANs) are trained to predict artifact-free reconstructions from artifact-corrupted CBCT images.

- Jo et al. (2025) quantitatively evaluated a proprietary deep-learning MAR algorithm built into a T2 Plus CBCT system.

- Using a standardized quality assurance phantom, they compared the performance of MAR under varying exposure protocols.

- Results showed a 61.5% reduction in artifact-affected areas under standard dose protocols, and up to 80.3% reduction in ultra-low-dose scans.

- Importantly, the study highlighted that AI-driven MAR maintained robust performance across different scanning conditions, indicating adaptability and clinical potential.

- Deep learning methods offer the benefit of nonlinear artifact suppression while preserving anatomical structures, but their effectiveness depends heavily on the quality and diversity of the training datasets.10,7

Clinical Perspective:

- Projection- and sinogram-domain approaches often preserve anatomical accuracy better, as they are rooted in physics-based corrections, but they may be limited when artifacts are very severe or when segmentation is inaccurate.

- Iterative reconstructions provide a balance between artifact suppression and preservation of detail, though computational demand can limit real-time application in clinical workflows.

- Deep learning methods show great promise in terms of efficiency and robustness but require extensive validation across scanners, exposure settings, and clinical conditions before widespread adoption.

- Across all methods, a critical consideration is that although artifact visibility is reduced, the true clinical accuracy—for tasks such as detecting peri-implant bone loss, identifying root fractures, or measuring bone density—must be independently verified.

Evaluation of MAR in Dental Applications:

Phantom and Quantitative Studies:

Several studies used phantoms to measure MAR efficacy. Bechara et al. (2012) used a phantom consisted of water-filled epoxy resin blocks with embedded metal beads to simulate bone next to high-density dental materials on a Vatech Picasso CBCT to compare MAR on vs off. They reported that MAR markedly suppressed streaks – for example, enabling MAR cut grey‐level variance by roughly 26–44 units across beam profiles and raised the contrast‐to‐noise ratio (CNR) by about +7.1 at 70 kVp. Higher kVp and centering the object further improved image quality.1

Jayakody et al. (2025) used dual‐energy (80/100/120 kV) CBCT with material decomposition and virtual monoenergetic imaging (200 keV) on six dental‐material phantoms (e.g. zirconia, high‐gold, cobalt‐chrome, amalgam). They found that dual‐energy virtual monoenergetic imaging alone and especially VMI combined with projection‐domain inpainting MAR gave the lowest artifact scores. In fact, VMI + MAR significantly outperformed a routine 100 kV scan with the vendor’s MAR (p<0.01) across all materials. Even dense, high-atomic-number restorations such as zirconia and gold, which typically generate extensive streaks, benefited from this combined method. The study highlights the clinical potential of integrating spectral CBCT with algorithmic correction, suggesting that future systems may surpass traditional hardware-based MAR solutions by tailoring image reconstruction to both material composition and projection geometry.6

Jo et al. (2025) tested a GAN/CNN‐based MAR on the standard SEDENTEXCT IQ phantom at 95 kV (0.2 mm voxels) under standard, low and ultra‐low‐dose modes. They quantified artifact area by thresholding and found that MAR cut artifact areas by ~61.5%, 73.6%, and 80.3% in the standard, low and ultra‐low‐dose modes, respectively; artifacts for vertical‐rod orientation were reduced by ~80.7% versus ~63.3% for horizontal rods.7,10

Kim et al. (2020) employed 3D‐printed mandibular phantoms with four common restorations (amalgam, gold crown, PFM, zirconia) and scanned them at 70 and 100 kVp with 0.2–0.3 mm3 voxels. Using automated edge counting, they found MAR reduced artifact edges by ~17.3–55.4% overall, with gold crowns showing the largest drop (p<0.05) and finer voxels (0.2 mm) consistently enhancing MAR effectiveness.11

Khosravifard et al (2021) conducted a controlled in-vitro evaluation, integrating clinically realistic bone surrogates with an objective, edge-based metric for the quantification of metal artifacts. Two bovine-rib blocks were prepared for this purpose: one was implanted with a titanium fixture, while the other contained a human root fitted with a stainless-steel intracanal post. These blocks were subsequently imaged within a wax mandibular arch at three distinct intra-arch positions.

The scan parameters included a fixed voxel size and tube current, with controlled variations in the field-of-view, peak kilovoltage (kVp), and the activation status of Metal Artifact Reduction. 36 CBCT scans were taken from each scan, three axial slices were preserved for subsequent analysis.

Multivariate analysis of the acquired data revealed that FOV size and MAR mode were the predominant factors influencing artifact presence. Specifically, larger FOVs consistently yielded a significantly greater number of edges compared to smaller FOVs. Furthermore, disabling MAR substantially increased edge counts for both types of metallic inclusions examined. Conversely, object position and the relatively modest kVp variations investigated did not exhibit consistent or statistically significant effects, and higher-order interactions proved negligible within the experimental setup. The authors acknowledged a limitation, noting that because edge detection also captures fine bony anatomical detail, comparisons were rigorously valid only within the same sample. Nevertheless, their findings robustly advocate for the utilization of smaller FOVs and the activation of vendor-provided MAR functionalities to mitigate streak artifacts in clinical cone-beam computed tomography protocols.5

MAR in Endodontics:

Costa et al. (2020) assessed two CBCT systems (OP300 Maxio and Picasso Trio) for detecting fractured endodontic instruments in mandibular molars. Their results showed no significant difference in diagnostic accuracy or image noise with MAR enabled. The MAR tool did not enhance the visibility of fractured files, whether canals were filled or unfilled, leading the authors to conclude that MAR is not beneficial for identifying broken instruments in endodontic practice.12

Bezerra et al. (2015), working with the Picasso Trio CBCT unit, examined MAR performance in diagnosing vertical root fractures in teeth containing metal posts. They found that enabling MAR did not improve diagnostic outcomes and, in fact, produced additional distortions that could obscure fine fracture lines. This suggested that MAR may not be appropriate when subtle endodontic features are the target of evaluation.

Almeshari et al. (2023) offered a more nuanced perspective by testing the Planmeca ProMax 3D system at three kVp levels (70, 80, and 90) with four MAR settings (off, low, mid, high). Their results showed that low MAR at 90 kVp significantly improved accuracy for VRF detection, yielding the highest sensitivity, specificity, and AUC. Conversely, mid and high MAR settings reduced accuracy, particularly at 70 and 90 kVp, while no significant differences were found at 80 kVp

These findings suggest that MAR is not universally reliable in endodontics. For fractured instruments and metal posts, MAR has shown little to no diagnostic value, and in some cases it may even hinder interpretation. The only consistent advantage noted was in Almeshari’s study, where a carefully selected combination (low MAR with higher kVp) improved VRF detection. Overall, the role of MAR in endodontics appears context- and parameter-dependent, and its indiscriminate use cannot be recommended.

MAR in Periodontics:

## Implant and Bone Defect Assessment:

In implantology, MAR efficacy has been mixed. Sheikhi et al. (2020) evaluated the impact of a MAR algorithm on detecting peri-implant fenestrations and dehiscences using CBCT and found that although MAR reduced the visual intensity of artifacts, it did not significantly enhance diagnostic accuracy. Detection of dehiscence was more reliable than fenestration, and observer agreement remained only moderate. MAR improves image appearance but has limited value in truly improving diagnostic performance for peri-implant bone defects.14

Salemi et al. (2021) placed 36 implants in bovine rib bone with and without fenestration/dehiscence defects, scanned them on Cranex 3D and ProMax 3D with/without MAR and measured diagnostic accuracy2. They reported that MAR actually reduced performance: the area under the ROC curve, sensitivity and specificity for defect detection were higher without MAR. In both CBCT systems, enabling MAR decreased the AUC and diagnostic accuracy for detecting fenestrations/dehiscences2. CBCT can detect peri-implant defects, the application of the MAR algorithm does not enhance detection2.

Nomier et al. (2022) conducted a similar in vitro study with 30 bovine blocks and simulated implant defects. They compared high-dose vs. low-dose CBCT with/without MAR (OP300 Maxio) and found a subtle effect. Without MAR, low-dose and high-dose had similar accuracy. When MAR was applied, high-dose imaging showed slightly higher AUC, but importantly the MAR tool decreased overall diagnostic ability, especially for smaller dehiscence. The application of MAR tool decreases the diagnostic ability of both defects, especially for dehiscence15. MAR did not improve detection of peri-implant fenestration or dehiscence, and even reduced sensitivity.

Multiple implant studies conclude that commercial MAR algorithms do not improve and can even degrade the detection of peri-implant bone defects. The visual clarity of the image may increase but critical diagnostic features are less effective.2,14,15

## Orthodontic bracket Artifacts Reduction:

Shavakhi et al., conducted a study in human skull with 20 brackets made of stainless steel attached to the maxillary dentition from second premolar to premolar placing buccal and palatal surfaces in two groups where palatally positioned brackets showed increased in contrast-to-noise ratio when MAR algorithms used and no significant changes with/without MAR algorithms in bracket and wire positions.16

# Summary of Findings:

Overall, MAR algorithms reduce visible artifacts but do not uniformly improve diagnostic metrics. Key observations include:

- Improved image metrics (phantoms): Quantitative analyses show MAR can markedly reduce artifact area or edge count in phantoms. Jo et al. reported >60% artifact-area reduction in a standardized phantom with deep-learning MAR10. In Bechara’s phantom test, enabling MAR increased contrast to noise ratio1. Artifact reduction is often material dependent (greater for high-attenuation materials like gold) and better at higher kVp or finer resolution1,11.

- Limited clinical gain: Salemi and Nomier both found that MAR reduced sensitivity and AUC for implant defect detection2,15. Thus, even though the quality of image was increased important details were lost.

- Algorithm dependence: The effectiveness of MAR varies by software and settings. Kim et al. noted that MAR effectiveness increased at smaller voxel sizes11. Bechara observed better MAR performance at higher kVp1.

- Potential trade-offs: MAR may introduce new artifacts or biases.

Bechara noted MAR altered grey values even in metal-free scans, raising concerns about wrongful alterations of anatomy1. The reduction of artifacts can sometimes come at the expense of blurring edges or creating false density changes. Although MAR algorithms enhance image clarity, they are not universally applicable solutions and care should be employed in clinical decisions based on the diagnostic requirements.

The collective evidence suggests that MAR tools in dental CBCT must be applied cautiously and case specifically. While vendor MAR algorithms can improve the visual quality of images and suppress obvious streaks1, they do not reliably improve and may worsen diagnostic detection of small defects2,15. Reasons include over-smoothing of fine structures and misclassification of true anatomy as artifact.

For tasks like fracture or defect detection, current MAR implementations appear to reduce diagnostic accuracy2,15. Clinicians should consider reviewing scans with MAR both on and off, especially when subtle pathology is suspected. High-resolution scanning (smaller voxels) can itself mitigate artifacts without MAR11.

Advances in MAR algorithms like adaptive sinogram weighting and deep learning show promise for stronger artifact removal. However, even these require thorough evaluation in patient images. Importantly, standardized testing protocols are needed. Jo et al. (2025) highlighted the value of a uniform phantom and metric for comparing MAR performance across studies10. Enabling MAR can inadvertently decrease the diagnostic ability for some defects. Until MAR methods are proven in vivo, use of these algorithms should be guided by the specific diagnostic question and results should be interpreted with caution and if possible correlated with non-MAR images.

Recent innovations in AI-assisted reconstruction methods have shown significant improvements in artifact suppression and diagnostic performance. Ge et al. (2024) proposed a comprehensive dual-energy method combining dual-energy acquisition, virtual monoenergetic projections, normalized MAR, and iterative reconstruction. Tested on phantoms with metallic restorations and implants, this approach significantly reduced streak artifacts and beam hardening while preserving fine structural details. Quantitative results showed improved similarity to reference images (SSIM rising to ~99.9%) and reduced error values (RMSE decreasing to ~71 HU), indicating clearer visualization of adjacent tissues without diagnostic information loss. The study highlights the potential of dual-energy MAR to outperform conventional techniques by enhancing image clarity and diagnostic reliability, though clinical validation remains necessary.17

Conclusion:

Metal artifact reduction (MAR) algorithms have shown significant advancements in enhancing image quality in dental CBCT, particularly in the presence of high-density restorative materials and implants. In tests of fractured tooth, fractured endodontic files and peri- implant defects MAR often failed to improve detection and sometimes reduce the diagnostic accuracy. Clinicians should therefore apply MAR tools selectively and interpret MAR processed images critically. Ongoing research especially using standardized evaluation methods is needed to optimized MAR for routine dental practice.

References:

- Bechara BB, Moore WS, McMahan CA, Noujeim M. Metal artefact reduction with cone beam CT: an in vitro study. Dentomaxillofacial Radiology. 2012 Mar 1;41(3):248-53.

- Salemi F, Jamalpour MR, Eskandarloo A, Tapak L, Rahimi N. Efficacy of metal artifact reduction algorithm of cone-beam computed tomography for detection of fenestration and dehiscence around dental implants. Journal of biomedical physics & engineering. 2021 Jun 1;11(3):305.

- Amirian M, Barco D, Herzig I, Schilling FP. Artifact reduction in 3D and 4D cone-beam computed tomography images with deep learning: a review. Ieee Access. 2024 Jan 12;12:10281- 95.

- Hegazy MA, Cho MH, Cho MH, Lee SY. Metal artifact reduction in dental CBCT images using direct sinogram correction combined with metal path-length weighting. Sensors. 2023 Jan 23;23(3):1288.

- Khosravifard A, Saberi BV, Khosravifard N, Motallebi S, Kajan ZD, Ghaffari ME. Application of an auto-edge counting method for quantification of metal artifacts in CBCT images: a multivariate analysis of object position, field of view size, tube voltage, and metal artifact reduction algorithm. Oral Surgery, Oral Medicine, Oral Pathology and Oral Radiology. 2021 Dec 1;132(6):735-43.

- Jayakody D, Agrawal H, Räinä E, Sipola A, Näpänkangas R, Ylisiurua S, Nieminen MT, Brix M. A Hybrid Approach Combining Dual-Energy and Inpainting Methods for Metal Artifact Reduction in Dentomaxillofacial CBCT: A Proof-of-Concept Phantom Study: D. Jayakody et al. Annals of Biomedical Engineering. 2025 Jul 20:1-0.

- Hegazy MA, Cho MH, Lee SY. Half-scan artifact correction using generative adversarial network for dental CT. Computers in Biology and Medicine. 2021 May 1;132:104313.

- Schulze R, Heil U, Groβ D, Bruellmann DD, Dranischnikow E, Schwanecke U, Schoemer E.Artefacts in CBCT: a review. Dentomaxillofacial Radiology. 2011 Jul 1;40(5):265-73.

- Hsieh J, Molthen RC, Dawson CA, Johnson RH. An iterative approach to the beam hardening correction in cone beam CT. Medical physics. 2000 Jan;27(1):23-9.

- Jo GD, Park CW, Jeon KJ, Han SS. Quantitative evaluation of metal artifact reduction in CBCT under varying exposure modes and rod orientations. Scientific Reports. 2025 Jul 1;15(1):20645.

- Kim YH, Lee C, Han SS, Jeon KJ, Choi YJ, Lee A. Quantitative analysis of metal artifact reduction using the auto-edge counting method in cone-beam computed tomography. Scientific Reports. 2020 Jun 1;10(1):8872.

- Costa ED, Brasil DM, Queiroz PM, Verner FS, Junqueira RB, Freitas DQ. Use of the metal artefact reduction tool in the identification of fractured endodontic instruments in cone‐beam computed tomography. International endodontic journal. 2020 Apr;53(4):506-12.

- Ruiz DC, Rosado LP, Fontenele RC, Farias-Gomes A, Freitas DQ. Vertical root fracture diagnosis in teeth with metallic posts: Impact of metal artifact reduction and sharpening filters. Imaging Science in Dentistry. 2024 Apr 2;54(2):139.

- Sheikhi M, Behfarnia P, Mostajabi M, Nasri N. The efficacy of metal artifact reduction (MAR) algorithm in cone‐beam computed tomography on the diagnostic accuracy of fenestration and dehiscence around dental implants. Journal of Periodontology. 2020 Feb;91(2):209-14.

- Nomier AS, Gaweesh YS, Taalab MR, El Sadat SA. Efficacy of low-dose cone beam computed tomography and metal artifact reduction tool for assessment of peri-implant bone defects: an in vitro study. BMC Oral Health. 2022 Dec 17;22(1):615.

- Shavakhi M, Soltani P, Aghababaee G, Patini R, Armogida NG, Spagnuolo G, Valletta A. A Quantitative Evaluation of the Effectiveness of the Metal Artifact Reduction Algorithm in Cone Beam Computed Tomographic Images with Stainless Steel Orthodontic Brackets and Arch Wires: An Ex Vivo Study. Diagnostics. 2024 Jan 10;14(2):159.

- Ge W, Liu Z, Cui H, Yuan X, Yang Y. A comprehensive dual energy method for CBCT metal artifact reduction. Physics in Medicine & Biology. 2024 Dec 27;70(1):015015.

- Hartshorne, J. (2022). Essential guidelines for using cone beam computed tomography (Cbct) in dentistry. Radiation dose, risks, safety, ethical and medico-legal considerations. Int Dent Afr Ed., 12, 5.

- Spagnuolo, G. (2022). Cone-beam computed tomography and the related scientific evidence. Applied Sciences, 12(14), 7140.

- Heller, Z. A., Hogge, M., Ragan, M. R., & Portnof, J. E. (2024). Applications of Cone Beam Computed Tomography Scans in Dental Medicine and Potential Medicolegal Issues. Dental Clinics, 68(1), 55-65.