Review Article

Compositional Changes For Reduction of Maximum Polymerization Temperature of PMMA Bone Cement For Anchoring Total Joint Arthroplasties: State-of-the-Art Review

ABSTRACT

Background: Poly(methyl methacrylate) (PMMA) bone cement is widely used in orthopaedics, particularly for anchoring an arthroplasty in the bone bed. Preparation of the cement involves mixing its powder and liquid, resulting in a dough that polymerizes, culminating in a hard product. The polymerization is an exothermic reaction, with maximum temperature reached (Tmax) being very high (in some cases, > 70 oC), depending on intrinsic variables, such as the cement composition, and extrinsic/process variables, such as powder-to-liquid ratio and the mixing method. It has been postulated that high Tmax contributes to or is responsible for thermal necrosis of periprosthetic tissue. Thus, methods to reduce Tmax have been a very active area of research. Purpose: The present work is a critical review of literature studies that involved one intrinsic variable, namely, change(s) to the composition of the cement, relative to that of the control cement. Methods: Using relevant search words, a search was conducted on appropriate databases, such as Google Scholar, PubMed, and Scopus, for English-language articles on methods to reduce Tmax and which also met other criteria. Results: 44 articles met the selection criteria and were reviewed in detail. It was found that 29 unique Tmax reduction methods have been used. 12 cement formulations yielded Tmax of < 44 oC, which, according to literature reports, is the lowest temperature at which thermal necrosis of periprosthetic tissue is likely to be initiated. Two examples of these formulations are 1) one in which 3 vol./vol.% of 1-dodecyl mercaptan was dissolved into the liquid of the control cement and 2) one in which 30 wt./wt.% of microcapsules containing paraffin/poly(methyl methacrylate–methylene bisacrylamide) (MPn) (n = ratio of MMA content to paraffin content = 2) were mixed with the powder of the control cement. The review includes a discussion of twelve shortcomings of the literature, which, therefore, indicated areas for future studies. These include determination of Tmax for a large number of commercial-produced and experimental formulations of antibiotic-loaded bone cement, development of a robust and validated methodology for calculating thermal necrosis potential, and investigation of the use of machine learning models to predict Tmax. The review ends with a compact summary of the main points made.

Keywords: Poly(methyl methacrylate) bone cement;  maximum polymerization temperature; thermal necrosis.

- INTRODUCTION

70 years after John Charnley reported the first use of poly(methyl methacrylate) (PMMA) bone cement                             (hereafter, “bone cement”) in preparing the grouting/anchoring/fixation bed for a primary hip implant [1, 2],  uses of this polymer in surgery have expanded to include a large assortment of applications, such as delivery of anesthetic agents [3], delivery of chemotherapeutic drugs [4], production of antibiotic-loaded  spacer for use in the first stage of a two-stage revision hip arthroplasty following serious or late chronic infection of the primary arthroplasty [5], treatment of bone defects or filling of a rare but locally aggressive neoplasm, such as giant cell tumor of bone [6,7], cranioplasty [8], restoration and/or augmentation of osteoporosis-induced compression fracture of vertebral body/bodies or treatment of pathologic spine compression fractures (for example, vertebroplasty (VP), balloon kyphoplasty (BKP), and vesselplasty [9-12]),                         and treatment of tinnitus [13]. However, the predominant use of bone cement remains in cemented primary and revision  arthroplasties [14-16].  In that respect, there is a very large number of plain and antibiotic-loaded brands that are approved by relevant national regulatory bodies, such as the United States Food and Drug Administration (USFDA) and the               United Kingdom Medicines and Healthcare Regulatory Agency [17-23]. For example,  in 2023, bone cement was used in ~85% of primary total hip arthroplasties (THAs) in Sweden [19], in ~96% of primary total knee arthroplasties in Germany [20], and, in the United States, there are those who recommend cementing the femoral component of a THA in some            sub-sets of patients, notably women who are >70 years old and those who are osteoporotic or have suffered displaced femoral neck fragility fracture(s) [24]. Additionally, there are reports on myriad experimental plain and antibiotic-loaded bone cement (ALBC) formulations and off-label ALBC formulations, with the latter being commonly used in clinical                              work [1, 2, 25-30].

There is a voluminous body of literature on the characterization of both approved brands and experimental formulations of bone cement [1, 2, 25-30], with the maximum exothermic temperature reached during mixing of the cement (Tmax) being a property that has often been the subject of the reports. Thus, Tmax as high as 124 oC have been reported [1-3] and the potential deleterious effect of a  high Tmax in a cemented arthroplasty have been reiterated, with examples being cause of or involvement in thermal necrosis of periprosthetic tissues, impaired blood circulation, and high predisposition to membrane formation at the cement-bone interface [1-3, 6]. Although there are many reviews of this body of  literature, only a few have included studies of methods to reduce Tmax [25, 31-33]. However, these reviews have limitations. Thus, the focus of the review by Berliner et al. [25] was prophylactic use of ALBC to treat/manage periprosthetic joint infection (PJI).                         Work on improving the mechanical properties of bone and imparting bioactivity were the main topics covered in the review by Seesala et al. [31]. A large proportion of the review by Szoradi et al. [32] was devoted to aspects such as the thickness and temperature of the cement mantle in an implant, aspects which relate to the structure (the cement mantle) rather than to the cement per se. In the review by He et al. [33],  only four studies of reduction of Tmax were discussed and one of these involved use of an extrinsic/process variable (herein, a variable that does not involve the compositional details of the cement formulation is designated an extrinsic/process variable, examples being the temperature at which the cement powder and/or liquid were stored prior to mixing, the powder-to-liquid ratio (PLR) used in mixing the cement, and the mixing method), as opposed to a compositional variable. Thus, shortcomings of reviews of the literature on bone cement with respect to reduction of Tmax persist.

The purpose of the present work was to present a comprehensive, robust, and critical review of the literature on use of modification of the composition of bone cement (an intrinsic variable) as a means to reduce its Tmax.                                                 The present  review is organized as follows. In Section 2, the methodology that was used to identify the studies that were included in the review is described. In Section 3, essential aspects of the determination of Tmax are given.                                      The methods that have been used to reduce Tmax are critically discussed in Section 4. The shortcomings of the literature (and, hence, areas for future studies) are highlighted in Section 5 and the main points made in the review are summarized in Section 6.

2.   METHODOLOGY

Over the period January 2025-September 2025, detailed searches were conducted in appropriate databases, such as Google Scholar, PubMed, Scopus, and Web of Science, for articles in English-language peer-review archival journals that contained both search words, namely, PMMA bone cement and polymerization temperature. These searches yielded              90 such publications of which 41 were eliminated from further consideration because they were not in peer-review          journals (23) or were available in abstracts only (18). Of the remaining  49 articles, 4 were excluded for a variety of reasons, namely, use of the cement was for delivery of a non-steroidal anti-inflammatory drug related to periodontitis [34],                        use of the cement was for VP or BKP [35], the control cement used was for dental application [36], the study was an investigation of the influence of  an extrinsic process variable (such as method of mixing the cement powder and liquid)         on Tmax [37], and part of one article was excluded because it was an investigation of the influence of an extrinsic/process              variable (PLR used in preparing the cement) on Tmax [38].  Thus, the present review was of the remaining 44 articles.

- BACKGROUND

With a few exceptions, determination of Tmax involved using a mold and a protocol specified in PMMA bone cement testing standards, such as ASTM F451 [39], ISO 5833 [40], and the Chinese National Standard YY0459-2003/ISO5833:20020 [41]. A variety of molds have been used [42-54], including  a cylindrical porcine trabecular bone mold (diameter and height =          9 mm and 12 mm, respectively) [42], a rectangular ultra-high-molecular-weight polyethylene mold (75 mm x 25 mm x 7 mm) [43] (Figure 1), and a short cylindrical poly(vinyl chloride) mold (diameter  and height =  25 mm and 40 mm, respectively) in which there was an Al foil to prevent cement dough leakage [51]. In one exception study, the prepared cement dough was manually shaped into a spherical blob, with radius of ~15 mm [54].

Figure 1.   Photograph of a typical experimental set-up, showing an ultra-high-molecular-weight polyethylene mold (left panel) and a digital thermometer (right panel) inserted into a port in the mold (right panel).

Equipment used in the determination of Tmax are a thermocouple whose welded end was used as a temperature sensor and, as such, was inserted into the center of a spherical cement dough blob [54] or a thermocouple/digital thermometer that is inserted into a port in the mold [43] (Figure 1) or directly into the test container [38].  When a mold was used, during the test, it was maintained at ambient temperature (23 ± 2 oC) [43, 54, 55] or it was immersed in an oven/bath whose temperature was kept at  37 oC [45, 46].

With regard to the test protocol, the cement powder and liquid are mixed either in a chamber open to the ambient atmosphere (hand mixing) or in a vacuum mixing chamber, typically at ~-75 kPa (vacuum mixing), until a dough is obtained. Within ~1 min of this happening, the dough is transferred from the mixing chamber to the plunger of the mold from which it is delivered into the mold, a weighted lid is placed on top of the mold, and, after a further 1 min or so, recording of the temperature of the curing cement dough (T) starts and is continued until the dough is fully cured,                                           with the T-versus-time (t) readings being recorded (usually, every 10 s or so) and exported to a data logger/data acquisition system for display and storage, using a software package. From the T-t plot, Tmax is determined (Figure 2).                       Typically, for a given cement formulation, 3 replicate tests were run.

Figure 2. A schematic presentation of test results.

4. METHODS USED TO REDUCE Tmax

4.1.  Prelude

Approved bone cement brands are formulations that are commercial-prepared, with the principal constituents of the liquid being a monomer (usually, methyl methacrylate (MMA)), a polymerization reaction accelerator/activator                                    (usually, dimethyl-p-toluidine (DMPT)), and a polymerization reaction retardation/stabilization agent (hydroquinone),              and the principal powder constituents being pre-polymerized PMMA beads with or without a PMMA-based co-polymer,              a radiopacifier (usually, particles of either BaSO4 or ZrO2), and a polymerization reaction initiator                                                   agent (benzoyl peroxide (BPO)). Methods used to reduce Tmax may be grouped into three categories, namely, those that target the composition of the liquid, the composition of the powder, or the composition of both the liquid and the                  powder [41-81].

4.2. Modifications to Composition of Liquid

The approaches taken were 1) partial substitution of MMA with a higher molecular weight methacrylate ester,                           such as isobornyl  or 4-N,N-(dimethylamino)phenethanol (QM-567) [56] or replacing DMPT with an alternative activator, namely, dicyclo-pentenyloxyethyl methacrylate (DEAPE) [56]; 2) substituting some of the MMA with 10-60 vol./vol.%                                                  of ethoxytriethyleneglycol methacrylate (TEG) [77]; 3) replacing DMPT with 4, N,N-dimethylaminobenzyl laurate, which is a long-chain acid derivative bearing an aromatic tertiary amine group [58]; 4) adding a methacrylic monomer derived from Vitamin E (MVE) to MMA (MVE content: 10-25 wt./wt.%) [47]; 5) in the case of CMWTM3 cement brand, substituting                 N,N-dimethylamino benzyl alcohol (DMOH) or 4,4’-dimethylamino benzdrol (BZN) for DMPT [59]; 6) introducing                         co-monomers, such as 4-methacryloyloxybenzoic acid (MBA), diethyl amino ethyl methacrylate (DEAEM), and                             4-diethylaminobenzyl methacrylate (DEABM)  (each in the range of 0.05-0.30 molar fraction (Mf)), to the MMA [46];                       7) using an alternative compound,  N,N-dimethylamino-4-benzyl laurate (DMAL) and N,N-dimethylamino-4-benzyl             oleate (DMAO), in place of DMPT [55]; 8) adding a chain-stopping agent (namely, 1-dodecyl mercaptan (DDM) [43, 54];                                9) mixing in multi-walled carbon nanotubes that either are unfunctionalized or carboxy-functionalized [60];                                    10) dispersing 0.5 wt./wt.% functionalized nano-TiO2 fibers (mean diameter and length = 12 nm and 2 m) [61];                       11) adding an N-alkyl imidazolium methacrylate monomer (MEIM-x)  (x = alkyl chain length = 6-18) with a total mass fraction of 2 wt.% [62]; 12) adding oxidized and silanized graphene powder, with two different methods being used for that           process [48]; 13) incorporating isobornyl acrylate (BA) [63]; and 14) adding polydimethylsiloxane (PDMS) [64].

4.3. Modifications to Composition of Powder

The approaches taken were 1) adding a mixture of hydroxyapatite (HA) and chitosan powder [42]; 2) in the case of a cement in which radiopacity was provided by BaSO4 particles, substituting particles of a ceramic that has a higher thermal conductivity (ZrO2 particles) [65]; 3) adding 10-30 vol./vol.% of short fibers (1-3 mm long) of sodium polyacrylate              (PAASf) [66]; 4) adding particles of a strontium-containing bioactive glass (SrBG) [67]; 5) mixing in microcapsules of a phase change material (paraffin or paraffin-based), which were synthesized using an emulsion polymerization method ([43, 44]; 6) adding chitosan (CS), silver nanowires (AgNWs), methacryloyl chitosan (CSMCC), or carboxy-functionalized multiwalled carbon nanotubes (FMWCNTs) [38, 68]; 7) mixing in poly(methyl methacrylate-co-acrylic acid) [P(MAA-AA)] microspheres [41]; 8) substituting BaSO4 for ZrO2 [52] and silanized ZrO2 for ZrO2 [52]; 9) adding one of  three variants of hydroxyapatite (HA), namely, plain hydroxyapatite (oHA), vinylmethoxysilane-treated HA (vHAP), or  PMMA-grafted HA  (gHAP) [49];            10) adding 7 wt./wt.% hallyosite (HNT) fillers [70]; 11) using PMMA nanoparticles synthesized using a mini-emulsion polymerization, with sodium dodecyl sulphate (SDS) as an anionic surfactant, virgin coconut oil (VO) as a co-stabilizer, potassium peroxodisulfate as the initiator, and aqua DM instead of PMMA [71]; and 12) adding a mixture of chitosan (CS) and mineralized allograft bone (human-derived) (BP) powders (with it consisting mainly of HA; particle size:                               150-1000 m) [72].

4.4. Modifications to Compositions of Powder and Liquid

A drop in Tmax was obtained through increase of the mean particle size (Dm) and the particle size distribution of the                    pre-polymerized PMMA beads in the cement powder or, for given combination of Dm and BPO content, decrease                          of the DMPT content in the liquid, or for a given combination of Dm and DMPT, decrease of  the BPO content in the            powder [45]. Among the aforementioned three effects, the greatest was produced by the increase in Dm. The rationale given was that during the polymerization process, the small PMMA beads (< 20 m) completely dissolved in MMA and only the large beads (>50-60 m) remained to release the heat produced during the polymerization process [45].

Careful variation of the concentration of DMPT and the amount of BPO led to a drop in Tmax [74].  Thus, for a given DMPT concentration (in the range 0.5-2 wt./wt.%), Tmax decreased monotonically with decrease in BPO amount, and, for a given BPO amount, Tmax decreased with a decrease in DMPT concentration although this trend was not  monotonic [74].                  From the results presented in this report [74], the present author obtained the following multilinear equation for the conjoint influence of DMPT content and BPO content on Tmax

Tmax (oC) = 54.61 + 12.4 (DMPT content) + 14.07 (BPO content)                   Adjusted R2 = 0.66

With a cement whose powder included pretreated 20 wt./wt.% calcium acetate [Ca(CH3COO)2] and whose liquid included Bis[2-methacryloyloxy)ethyl] phosphate (B2meP), Tmax dropped monotonically with increase in B2meP content           (between 0 and 30 vol./vol.%) [69].

4.5. Overview

In an effort to reduce Tmax, about equal research attention has been paid to modifications of the composition of the liquid of the control cement as to modifications of the composition of its powder. Across the compositional changes reported in the studies discussed above, the resulting  % decrease of Tmax obtained ranged from 1.4 to 56.9, with the mean, median, and population standard deviation being 20.4, 24.6, and 14.4, respectively. When the index of acceptable improvement in reducing Tmax is taken to be ≥50 % decrease, 7 formulations meet this criterion, 4 of which involved modification of the composition of the liquid (Table 1). However, this approach measures effectiveness of  Tmax reduction method rather than the potential  clinical efficacy of an experimental formulation. For that insight, the results should be discussed                               vis a vis  the postulated  criterion for initiation of thermal necrosis of periprosthetic tissue. To do this, however, requires overcoming two challenges. First,  in some studies, this criterion has been given as a threshold temperature (Tinc) only whereas, in other studies, it has been given in terms of Tinc and its duration. Second, various values of Tinc have been presented, these being 44-47 oC [82], > 55 oC [83], >56 oC [84], > 70 oC [85], and > 90 oC  [75]. Since, in nearly all of the studies reviewed in the present work, both temperature and time at that temperature results were not given, an evaluation of the potential clinical usefulness of the results is limited to reported Tmax when compared to Tinc. When this is done using the lowest value of Tinc postulated in the literature (44 oC), it was found that 12 experimental cement formulations have potential for use in anchoring arthroplasties (Table 2).

| Table 1. Summary of experimental cement formulations that produced % decrease in Tmax of ≥ 50% (relative to the mean value for the control cement) |  |  |  |
| --- | --- | --- | --- |
| Experimental cement formulation identifier | Compositional informationa | % decrease | Source of results [Ref. #] |
| 0.30-DEAEM | Liquid formulation includes 0.30 molar fraction of a co-monomer, DEAEM | 56.9 | Cervantes-Uc et al. [46] |
| PMMA-BC-60%AA | 60% molar ratio of P(MMA-AA) microspheres were mixed with the powder of the control cement. | 55.8 | Chen et al. [41] |
| 0.20-DEAEM | Liquid formulation includes 0.30 molar fraction of a co-monomer, DEAEM | 54.5 | Cervantes-Uc et al. [46] |
| 73PMMA-PAASf | PAASf were blended with the powder of the control cement (proportion of PMMA to PAASf = 7:3) | 54.4 | Tang et al. [66] |
| PMMA-BC-50%AA | 50% molar ratio of P(MMA-AA) microspheres were blended with the powder of the control cement. | 54.0 | Chen et al. [41] |
| 25MVE | MVE was added to the liquid of the control cement in the amount of 25 wt./wt.% | 53.8 | Méndez et al. [47] |
| 5-MEIM-12 | 5 vol./vol.% of a quaternary ammonium monomer with N-alkyl chain length of 12 (MEIM-12) was dissolved into the MMA in the liquid of the control cement. | 51.8 | Zhu et al. [62] |
| aDetailed information on the compositions of the formulations are given in the text. |  |  |  |

| Table 2. Summary of information about experimental cement formulations with Tmax ≤ 44 ℃ |  |  |  |
| --- | --- | --- | --- |
| Experimental cement formulation identifier | Compositional informationa | Tmax (℃) | Source [Ref. #] |
| EXP-60TEG | In the liquid, 60 vol./vol.% of TEG was blended with the MMA | 44 | Pascual et al. [77] |
| 5-MEIM-10 | 5 vol./vol. of a quaternary ammonium monomer with N-alkyl chain length of 10 (MEIM-10) was dissolved into the MMA in the liquid of the control cement. | 42.2 ± 1.6 | Zhu et al. [62] |
| 5-MEIM-12 | 5 vol./vol. of a quaternary ammonium monomer with N-alkyl chain length of 12 (MEIM-12) was dissolved into the MMA in the liquid of the control cement. | 40.4 ± 3.3 | Zhu et al. [62] |
| 0.20-DEAEM | Liquid formulation includes 0.20 molar fraction of DEAEM as co-monomer. | 40.1 ± 1.0 | Cervantes-Uc et al. [46] |
| 2DDM | 2 vol./vol.% of DDM was dissolved into the liquid of the control cement. | 39.4 ± 1.7 | McKee et al. [43] |
| 0.30-DEAEM | Liquid formulation includes 0.30 molar fraction of DEAEM as co-monomer. | 38.0 ± 1.0 | Cervantes-Uc et al. [46] |
| BC-20MPn1 | 20 wt./wt.% of microcapsules containing MPn (n = ratio of MMA to paraffin = 1 ) were blended with the powder of the control cement | 37.6 | Xia et al. [44] |
| 25MVE | 25 wt./wt.% of MVE was added to the liquid of the control cement. | 35.8 ± 0.0 | Mendez et al. [47] |
| 15MEPAR | 15 wt./wt.% paraffin microcapsules were blended with the powder of the control cement. | 35.2 ± 2.1 | McKee et al. [43] |
| 3DDM | 3 vol./vol.% DDM was dissolved into the liquid of the control cement. | 34.2 ± 1.4 | McKee et al. [43] |
| 25MEPAR | 25 wt./wt.% paraffin microcapsules were blended with the powder of the control cement. | 32.4 ± 2.3 | McKee et al. [43] |
| BC-30MPn2 | 30 wt./wt.% of microcapsules containing (MPn) (n = ratio of MMA to paraffin = 2) were blended with the powder of the control cement | 32.1 | Xia et al. [44] |
| aDetailed information on the compositions of the formulations are given in the text. |  |  |  |

5. SHORTCOMINGS OF THE LITERATURE AND AREAS FOR FUTURE STUDY

Several shortcomings of the studies covered in this review are noted, which suggest twelve areas for future study.

First, in some studies, the experimental cement formulations were radiolucent [45, 67, 69, 71, 76, 77, 78],                                           a control cement formulation was not included in the formulations evaluated [45, 71, 74], one reference cement                  (herein, defined as an approved or commercial-prepared brand) was not included [41, 46, 47, 48, 49, 52, 53, 54, 62, 65, 67, 69, 74, 77, 78], or the reference cement brand was taken to be the control cement [38, 50, 68]. In each future study,                      a separate control cement and at least one reference cement brand should be included. Additionally, future studies should include novel experimental formulations for which attractive properties have been reported but whose Tmax was not determined [68, 75, 78, 86-88].  An example is a cement that contains two-dimensional magnesium phosphate nanosheets and hydroxyapatite nanofibers, whose attractive properties include enhanced quasi-static compressive strength and excellent bioactivity [68].

Second, the method used to blend an additive to the liquid of a cement formulation or to mix/combine additive(s) to the powder of a cement formulation was stated in only a few reports. In terms of the liquid, the method used was ultrasonic agitation [61]. As for the powder, the devices used include a vortex mixer [38] and “a three-dimensional motion mixer” [67]. In all future studies, the same mixing methods should be used so that when study results are compared, mixing method is not a confounding variable.

Third, there are three deficiencies that appear in many studies. One, at least one commercial-prepared and approved cement brand was not among those whose Tmax was determined. Two, either the PLR used in preparing an experimental formulation was not stated [56, 66] or a value other than 2 g mL-1, (which is recommended by companies for mixing nearly all commercial-prepared and approved cement brands; RPLR) was used [47,49, 52, 54, 57, 59, 60, 71, 76, 78].                   Three, with the exception of four studies [43, 50, 60, 70], the method used to mix the powder and the liquid was not             stated [48, 56, 74] or was hand mixing [46, 52, 54, 78, 89]. In current arthroplasty practice, a third-generation cementation technique is used, which entails mixing the powder and the liquid and delivering the resulting dough to the prepared bone bed using a vacuum mixing system [90]. Each of the above-mentioned  deficiencies should be considered when designing future studies.

Fourth, in one study (Kim et al. [50]), even though the work was on bone cements for use in vertebroplasty,                                     the reference cement that was used (CMWTM3) was, in fact, for anchoring arthroplasties. Thus, in future studies,                        this mismatch must be avoided.

Fifth, in nearly all the studies, the test was carried out at ambient laboratory temperature  (22-23 oC) rather than at 37 oC  (normal body temperature). It is known that handling properties of a bone cement are strongly affected by the temperature at which they are determined  [58]. Thus, there is opportunity for either developing and validating a new standard for the test or revising the existing standard to include conducting it at 37 oC. With regard to the former issue, alternative means of obtaining the temperature of the polymerizing dough should be included, an example being use of an infrared thermography camera [91].

Sixth, in many studies, the influence of two or more compositional variables on Tmax was investigated [45, 53, 56, 73, 74, 78-81]. For example, in the studies by Ginebra et al. [73] and Milner et al. [74], the combinations were amounts of BPO and DMPT and in the study by Girijia et al. [78], the combinations were amounts of a bioactive agent (Ca2MgSi2O7) and a radiopacifier (ZrO2) in the cement powder. In each of the above-referenced studies, the conjoint influence of the combination of variables used was not determined.  This lacuna presents an opportunity to perform an optimization study that would take the form of determining the optimum values of the variables to obtain a stated value of Tmax.                                    In this regard, the study by Ruiz Rojas et al. [53] should be instructive.

Seventh, in only two studies was an ALBC used as the control or reference cement [38, 43]. ALBCs are widely used in both primary and revision cemented arthroplasties for preventing and managing periprosthetic joint infection (PJI) [92-97].                Future in vitro characterization studies of ALBCs should include determination of Tmax and methods of reducing it, if needed.                         In each study, the roster of cements studied should include at least one commercial-prepared and approved ALBC brand, such as Palacos®R+G and Surgical SimplexTM with Tobramycin.

The current generation of approved ALBC brands (those in which the antibiotic is gentamicin, clindamycin, daptomycin, tobramycin, or vancomycin, or a combination  thereof) has many shortcomings, among which is development of resistance by bacteria that commonly present in PJI cases (such as methicillin-resistant S. aureus and S. epidermidis) to the           antibiotic [98, 99]. Thus, there is ongoing work to formulate and characterize novel antibacterial/antimicrobial/antibiotic-loaded PMMA bone cements (ABMALBCs) [100-109], such as 1) long-release type (in which an antimicrobial agent is released over an extended period of time; typically, > 10 weeks) [101, 102] and 2) one that contains a sortase A                inhibitor [107]. The eighth shortcoming is a shortage of studies that report Tmax of emerging ABMALBCs.                                 Future studies should address this.

Ninth, in several literature reports, a statistical analysis of the study results was not included [44, 45, 46, 54, 65, 66, 71, 74, 77, 80, 81, 89] and, in the studies in which this was done, there are issues with the method used. Thus, with the exception of three studies [43, 52, 70], the method(s) used were parametric methods, most commonly one-way analysis-of-variance (ANOVA) [34, 38, 41, 47, 53, 59, 62, 68, 82]. It is highlighted that a parametric method should be used only if it is first shown, through the use of a normality test (such as the Smirnoff-Kolmogorov test), that each of the datasets being compared is a normal distribution and through, for example Levene’s test, that there is homogeneity of variance between the datasets. If that is not the case, then, a non-parametric test of comparison should be used, an example being the Kruskal-Wallis test. In the interest of efficiency, in all future studies, non-parametric method(s) should be used.

Tenth, each of the approaches that have been taken, to date, in discussing the clinical relevance of Tmax suffers from twin shortcomings of being empirical and lacking validation. Future studies should develop a rigorous, first-principles-based, and validated methodology that yields a critical value for initiation of thermal necrosis in the periprosthetic tissue (TNINC).          Work in developing TNINC should be informed by results of studies on parameters that influence the T-versus-t behavior of a bone cement, such as, in biomechanical tests, it has been found that the extent of thermal necrosis at the cement-bone interface is influenced by implant material and cement mantle thickness [110]. Most importantly, it should be ensured that the methodology for determining TNINC is not computationally intensive and that the resulting TNINC is validated using results from, at the minimum, well-designed animal model studies.

Eleventh, there is lack of use of artificial intelligence to extend and enhance studies in this area. Thus, future studies should include use of machine learning (ML) to design bone cement formulations to achieve a value < TNINC (as developed in future studies discussed above). The foundational work in such an ML study would entail a comparison of the predictive performance of a slate of ML algorithms/models (such as extreme gradient boosting, k-nearest neighbor, and random forest) for predicting Tmax, as a function of a large collection of compositional variables, against experimentally-obtained values        of Tmax. For this comparison, the performance metrics should include coefficient of determination, mean absolute error,        and simulation execution time.

Twelfth, none of the experimental formulations that show particular promise in in vitro tests in studies reviewed in this work has been the subject of animal model studies. These studies are crucial in the path to evaluation in randomized controlled trials (RCTs), obtaining a patent and, eventually, approval, by regulatory bodies, for commercial production and, finally, clinical use. Ideally, for the RCT, the predicate cement must be a bone cement brand that is widely used in many countries in the world (for example, Simplex P® (plain variety) and Palacos R with Gentamicin® (an ALBC));                                                                    the total number of patients should be at least 1,000, and assigned randomly to the two study groups                                  (predicate cement and experimental bone cement formulation); the patients should be matched for relevant characteristics, such as age, body mass index, ratio of women to men, and incidence of pathology that necessitated arthroplasty                       (for example, end-stage osteoarthritis), and implant design; the RCT should be carried out at a variety of medical centers (such as large teaching hospitals, specialized surgery centers, and rural hospitals) located in many countries evenly spread across the world; and the minimum follow-up duration should be 5 years.

6. CONCLUSION

The following is a summary of the key points made in this review:

- PMMA bone cement is widely used in orthopaedic surgery, notably in cemented arthroplasties, where it is used to anchor the implant in the bone bed. Nonetheless, the cement many shortcomings, one of which is that its curing of the cement dough (mixture of powder and liquid) is an exothermic reaction, with the maximum temperature (Tmax)                   high enough that it may cause or contribute to thermal necrosis of periprosthetic tissue (with some cement formulations in current clinical use).

- A large body of literature exists on methods to reduce Tmax through modifications to the composition of the liquid or         of the powder or both the powder and the liquid. In the present work, studies involving 29 such methods were reviewed.                                     When  decrease of Tmax of  ≥50%   (relative to Tmax of the control cement formulation) is used as a criterion of the effectiveness of the methods, the most effective method was one in which some of the methyl methacrylate                         (the monomer in the control cement) was replaced with 0.30 mole fraction of diethyl amino ethyl methacrylate                  (Tmax decrease = 56.9%). When Tmax ≤44 oC (a temperature above which, it has been postulated, thermal necrosis of periprosthetic tissue is likely to occur) is used as the criterion of performance of the Tmax reduction methods reported, 12 experimental cement formulations are identified as meeting this criterion, with the lowest Tmax (32.1 oC) obtained using a cement in which 30 wt./wt.% of microcapsules containing paraffin/poly(methyl methacrylate–methylene bisacrylamide) (MPn) (n = ratio of MMA content to paraffin content = 2) were mixed with the powder of the control cement.

- The literature has several shortcomings, among which are paucity of studies that involve antibiotic-loaded PMMA bone cements (which are very widely used in revision hip and knee arthroplasties), lack of a robust relationship between Tmax of a cement and the temperature at which thermal necrosis will be initiated (Tnc), and use of machine learning to design formulations whose Tmax would be substantially lower than Tnc. These shortcomings should be addressed in future studies.

References

- Lewis G. Properties of acrylic bone cement: state of the art review. J Biomed Mater Res (Appl Biomater) 1997;                38: 155-182. https://doi.org/10.1002/(sici)1097-4636(199722)38:2%3C155::aid-jbm10%3E3.0.co;2-c.

- Kühn K-D. Bone cements. Springer Medizin, Berlin, Germany; 2014.

- Montfort-Mira M, Balaguer-Castro M, Jornet-Gibert M, Cito S, Flor-López S, Torner P. Mechanical integrity assessment of polymethylmethacrylate bone cement incorporating local anesthetics: implications for joint arthroplasty. J Arthroplasty 2025;       40: 2160-2166.  https://doi.org/10.1016/j.arth.2025.01.014.

- Phull SS, Yazdi AR, Ghert M, Towler MR. Bone cement as a local chemotherapeutic drug delivery carrier in orthopedic oncology: a review. J Bone Oncol 2021; 26: 100345. https://doi.org/10.1016/j.jbo.2020.100345.

- Batailler C, Cance N, Lustig S. Spacers in two-stage strategy for periprosthetic infection. Orthopaedics & Traumatology:      Surgery & Research 2025; 111: 104074. https://doi.org/10.1016/j.otsr.2024.104074.

- Mohammadi S-S, Phull SS, Bal BS, Towler MR. Bone void filler materials for augmentation in comminuted fractures:                                 a comprehensive review. J Orthop Surg Res 2025; 20: 449. https://doi.org/10.1186/s13018-025-05857-2.

- Costello II JP, Travis LM, Jahn K, Pretell-Mazzini JA. The role of bone grafting vs. bone   cement in the treatment of giant cell tumor of bone. JBJS Reviews 2024;12:e24.00080.
 https://doi.org/10.2106/JBJS.RVW.24.00080.

- Neill R, Harris P, Daggubati LH. Bone cement versus bone flap replacement: a comparative meta-analysis of posterior fossa craniotomy complications. Surg Neurol Int 2025; 16:25. https://doi.org/10.25259/sni_789_2024.

- Cine HS, Uysal E, Aladdam M, Herdan E, Gunaydin ME, Demir H, Karaarslan N. Comparative efficacy of unilateral vs. bilateral approaches in percutaneous kyphoplasty and vertebroplasty for osteoporotic vertebral compression fractures. Indian Journal of Neurosurgery 2025; 14: 052–058. https://doi.org/10.1055/s-0044-1800829.

- Aduri TT, Dhillon M, Bansal P, Vatkar A, Dhatt SS, Kumar V. Comprehensive systematic-review and meta-analysis: Treatment outcomes of unipedicular vs bipedicular approaches in vertebroplasty and kyphoplasty for osteoporotic vertebral compression fractures. J Clin Orthop Trauma 2025; 66:103010. doi: 10.1016/j.jcot.2025.103010.

- Agrawal P, Banode P, Nagendra V, Joshi R, Agrawal S, John J, Nawale N. Percutaneous vesselplasty for the treatment of vertebral osteoporotic collapse fractures. Cureus 2024; 9: 68695. https://doi.org/10.7759/cureus.68695.

- Wang K, Chen H, Feng F, Zhang S. Comparison of clinical efficacy between vesselplasty and percutaneous vertebroplasty in genant grade 3 osteoporotic vertebral compression fracture. World Neurosurg 2025; 196: 123751. https://doi.org/10.1016/j.wneu.2025.123751.

- Almutairi HM, Alkhalifah KM, Alhujaili HN, AlAmry S. Novel technique utilizing polymethylmethacrylate cement for the treatment of pulsatile tinnitus caused by different sigmoid sinus pathologies. Eur Arch Otorhinolaryngol 2025; 282: 1095-1101. https://doi.org/10.1007/s00405-024-08974-6.

- Auran R, Movassaghi K, Nam D, Heckmann N. Bone cement in adult hip and knee reconstruction: a review of commercially available options and clinical outcomes. J Am Acad Orthop Surg 2024; 32: 1057-1066. https://doi.org/10.5435/jaaos-d-23-01232.

- Alagha MA, Cobb J, Liddle A, Malchau H, Rolfson O, Mohaddes M. Adverse outcomes after cemented and cementless primary elective total hip arthroplasty in 60,064 matched patients: a study of data from the Swedish Arthroplasty register. J Arthroplasty 2025; 40: 1738-1744. https://doi.org/10.1016/j.arth.2025.01.003.

- Fontanellas-Fes A, Hinarejos P, Pérez-Prieto D, Martínez-Lozano J, Sánchez-Soler J, Torres-Claramunt R, Perelli S, Monllau JC. Comparable bone penetration between antibiotic-loaded and plain bone cement in total knee arthroplasty. Knee Surg Sports Traumatol Arthrosc 2025; 33: 364-370. https://doi.org/10.1002/ksa.12379.

- Carender CN, Hegde V, Levine BR, Huddleston JI, Cohen-Rosenblum A. Highlights of the 2024 American Joint Replacement Registry Annual Report. Arthroplasty Today 2025; 33: 101727. https://doi.org/10.1016/j.artd.2025.101727.

- Shum BCP, Lau LCM, Cheung JPY, Choi SW. Current status of Asian joint registries: a review. EFORT Open Rev 2025; 10: 250–257. https://doi.org/10.1530/EOR-2024-0085.

- W-Dahl A, Kärrholm J, Rogmark C, Johansson O, Arani PI, Mohaddes M, Rolfson O. Annual Report 2024.                       The Swedish Arthroplasty Register 2024.

- Humez M, Kötter K, Skripitz R, Kühn KD. Evidence for cemented TKA and THA based on a comparison of international register data. Orthopadie (Heidelb) 2024; 53: 597-607. https://doi.org/10.1007/s00132-024-04489-4.

- Perez-Prieto D, Koetter K, Fontanellas-Fes A, Martínez-Cruz O. Total knee arthroplasty revision risks depending on the bone cement used-data from 50,545 knee replacements of the Catalan Arthroplasty Registry. Journal of Experimental Orthopaedics 2025; 12: 70271. https://doi.org/10.1002/jeo2.70271.

- Gaidici T, Deckey DG, Seyler TM, Bolognesi MP, Spangehl MJ, Bingham JS. Femoral fixation for primary total hip arthroplasty-an international registry perspective. Arthroplasty Today 2025; 34: 101755. https://doi.org/10.1016/j.artd.2025.101755.

- Lauck BJ, Larabee L, Morningstar JL, Fraser JF. Global trends in fixation methods for primary total knee arthroplasty: An international registry-based analysis and review of outcomes. Journal of Orthopaedics 2025; 67: 230-237. https://doi.org/10.1016/j.jor.2025.05.031.

- Khanuja HS, Mekkawy KL, MacMahon A, McDaniel CM, Allen DA, Moskal JT. Revisiting cemented femoral fixation in hip arthroplasty. J Bone Joint Surg Am 2022; 104: 1024-1033. doi:10.2106/JBJS.21.00853.

- Berliner ZP, MD1, David Freccero D, Cooper HJ, Smith EL, Kain MS. Antibiotic loaded bone cement in primary total hip and knee arthroplasty. Montefiore Einstein Journal of Musculoskeletal Medicine and Surgery 2025; 4: 1-6. https://doi.org/10.1177/30498929251329316.

- Feng B, Wei Z, Zhu W, Xu Y, Li Z, Weng X. Mineralized collagen-enhanced bone cement: a prospective, randomized controlled trial in total knee arthroplasty. J Mater Sci Mater Med 2025; 36: 41. https://doi.org/10.1007/s10856-025-06893-4.

- Lewis G. Properties of antibiotic-loaded acrylic bone cements for use in cemented arthroplasties: a state-of-the-art review. J Biomed Mater Res B Appl Biomater 2009; 89: 558-574. https://doi.org/10.1002/jbm.b.31220.

- Lewis G. Alternative acrylic bone cement formulations for cemented arthroplasties: Present status, key issues, and future prospects. Journal of Biomedical Materials Research Part B: Applied Biomaterials 2008; 84B: 301-319. https://doi.org/10.1002/jbm.b.30873.

- Shirbacheh AK, Shirbache K. PMMA bone cement: properties, applications, and innovations. Journal of Composites and Compounds 2023; 5. https://doi.org/10.61186/jcc.5.4.6.

- Patel R, Mcconaghie G, Webb J, Laing G, Roach R, Banerjee R. An overview of bone cement: perioperative considerations, complications, outcomes and future implications. J Perioper Pract 2024; 34: 106-111. https://doi.org/10.1177/17504589231169861.

- Seesala VS, Sheikh L, Basu B, Mukherjee S. Mechanical and bioactive properties of PMMA bone cement:                            a review. ACS Biomater Sci Eng 2024; 10: 5939-5959. https://doi.org/10.1021/acsbiomaterials.4c00779.

- Szoradi GT, Feier AM, Zuh SG, Russu OM, Pop TS. Polymethyl methacrylate bone cement polymerization induced thermal necrosis at the cement–bone interface: a narrative review. Applied Sciences 2024; 14: 11651. https://doi.org/10.3390/app142411651.

- He H, Liu Y, Dong Y, Zhao W, Wang K, Guo K, Jiao E, Xu S, Ju G, Wang P, Ma D, Shi G, Zhang H. Material properties and progress in modification of hydrogel-based self-expandable poly (methyl methacrylate) bone cement. RSC Advances 2025; 15: 26959-26980. https://doi.org/10.1039/d5ra00592b.

- Lopes PP, Silva MS, Fernandes MH. Influence of ibuprofen addition on the properties of a bioactive bone cement.           J Mater Sci Mater Med 2013; 24: 2067-2076. https://doi.org/10.1007/s10856-013-4958-6.

- Lv Y, Li A, Zhou F, Pan X, Liang F, Qu X, Qiu D, Yang Z. A novel composite PMMA-based bone cement with reduced potential for thermal necrosis. ACS Appl Mater Interfaces 2015; 7: 11280-11285. https://doi.org/10.1021/acsami.5b01447.

- Król K, Macherzyńska B, Pielichowska K. Acrylic bone cements modified with poly (ethylene glycol)-based biocompatible phase-change materials. Journal of Applied Polymer Science 2016; 133. https://doi.org/10.1002/app.43898.

- Liu CW; Chou WH. Influence of bone cement mixing mechanisms on orthopedic surgery under various conditions and parameters. Journal of Biomaterials and Tissue Engineering 2017; 7: 863-867. https://doi.org/10.1166/jbt.2017.1639.

- De Mori A, Di Gregorio E, Kao AP, Tozzi G, Barbu E, Sanghani-Kerai A, Draheim RR, Roldo M. Antibacterial PMMA composite cements with tunable thermal and mechanical properties. ACS Omega 2019; 4: 19664-19675. https://doi.org/10.1021/acsomega.9b02290.

- ASTM F451. Standard Specification for Acrylic Bone Cement. American Society for Testing and Materials,                    West Conshohocken, PA, USA; 2021.

- ISO 5833. Implants for Surgery-Acrylic Bone Cement. International Organization for Standardization, Geneva, Switzerland, 2002.

- Chen L, Tang Y, Zhao K, Liu J, Bai H, Wu Z. Absorptive and expansive behaviors of poly(methyl methacrylate-co-acrylic acid) bone cement. Colloids Surf B Biointerfaces 2020; 189: 110848. https://doi.org/10.1016/j.colsurfb.2020.110848.

- Kim SB, Kim YJ, Yoon TL, Park SA, Cho IH, Kim EJ, Kim IA, Shin JW. The characteristics of a hydroxyapatite-chitosan-PMMA bone cement. Biomaterials 2004; 25: 5715-5723. https://doi.org/10.1016/j.biomaterials.2004.01.022.

- McKee R, Harris M, Hadley J, Zhang J, Lewis G. Comparative influence of two compositional modifications on maximum exotherm temperature and other properties of an antibiotic-loaded PMMA bone cement. Current Applied Polymer Science 2018; 2: 76-88. http://dx.doi.org/10.2174/2452271602666180727112518.

- Xia X, Shi R, Huang J, Li Y, Zuo Y, Li J. Development of a phase change microcapsule to reduce the setting temperature of PMMA bone cement. J Appl Biomater Funct Mater 2020;18:1-6. https://doi.org/10.1177/2280800020940279.

- Pascual B, Vázquez B, Gurruchaga M, Goñi I, Ginebra MP, Gil FJ, Planell JA, Levenfeld B, Román JS.                             New aspects of the effect of size and size distribution on the setting parameters and mechanical properties of acrylic bone cements. Biomaterials 1996; 17: 509-516. https://doi.org/10.1016/0142-9612(96)82725-6.

- Cervantes-Uc JM, Vázquez-Torres H, Cauich-Rodríguez JV, Vázquez-Lasa B, del Barrio JS. Comparative study on the properties of acrylic bone cements prepared with either aliphatic or aromatic functionalized methacrylates. Biomaterials. 2005; 26: 4063-4072. https://doi.org/10.1016/j.biomaterials.2004.10.030.

- Méndez JA, Aguilar MR, Abraham GA, Vázquez B, Dalby M, Silvio LD, San Román JS. New acrylic bone cements conjugated to vitamin E: curing parameters, properties, and biocompatibility. J Biomed Mater Res 2002;                           62: 299-307. https://doi.org/10.1002/jbm.10296.

- Paz E, Ballesteros Y, Abenojar J, Dunne N, Del Real JC. Advanced G-MPS-PMMA bone cements: influence of graphene silanisation on fatigue performance, thermal properties and biocompatibility. Nanomaterials (Basel) 2021; 11: 139. https://doi.org/10.3390/nano11010139.

- Tham DQ, Huynh MD, Linh NTD, et al. Pmma bone cements modified with silane-treated and PMMA-Grafted hydroxyapatite nanocrystals: preparation and characterization. Polymers (Basel). 2021; 13: 3860. https://doi.org/10.3390/polym13223860.

- Kim MJ, Park SY, Kang S, et al. Low modulus PMMA-based bone cement for the reduction of adjacent vertebral fractures after vertebroplasty. Acta Biomater 2025; 203: 399-411. https://doi.org/10.1016/j.actbio.2025.07.053.

- Stoops K, Brown JM, Santoni B, Groundland J. Thermal properties of polymethyl methacrylate vary depending on brand and type. J Orthop Res 2023; 41: 614-618. https://doi.org/10.1002/jor.25389.

- Ayre WN, Scully N, Elford C, Evans BA, Rowe W, Rowlands J, Mitha R, Malpas P, Manti P, Holt C, Morgan-Jones R, Birchall JC, Denyer SP, Evans SL. Alternative radiopacifiers for polymethyl methacrylate bone cements:                   Silane-treated anatase titanium dioxide and yttria-stabilised zirconium dioxide. J Biomater Appl 2021; 35: 1235-1252. https://doi.org/10.1177/0885328220983797.

- Ruiz Rojas LM, Valencia Zapata ME, Gordillo Suarez M, Advincula R, Grande-Tovar CD, Mina Hernández JH. Optimization of mechanical and setting properties in acrylic bone cements added with graphene oxide.  Applied Sciences 2021; 11: 5185. https://doi.org/10.3390/app11115185.

- Endogan T, Şerbetçi K, Hasirci N. Effects of ingredients on thermal and mechanical properties of acrylic bone cements. Journal of Applied Polymer Science 2009; 113: 4077 – 4084. http://dx.doi.org/10.1002/app.30488.

- Lewis G, Xu J, Deb S, Lasa BV, Román JS. Influence of the activator in an acrylic bone cement on an array of cement properties. J Biomed Mater Res A 2007; 81: 544-553. https://doi.org/10.1002/jbm.a.31067.

- Brauer GM, Steinberger DR, Stansbury JW. Dependence of curing time, peak temperature, and mechanical properties on the composition of bone cement. J Biomed Mater Res 1986 20: 839-852. https://doi.org/10.1002/jbm.820200614.

- Pascual B, Gurruchaga M, Goni I, Ginebra MP, Gil FJ, Planell JA, Levenfled B, Vázquez B, Román JS. Mechanical properties of a modified acrylic bone cement with etoxytriethyleneglycol monomethacrylate.  J Mater Sci: Mater Med 1995; 6: 793-798 https://doi.org/10.1007/BF00134319.

- Vazquez B, San Roman J, Deb S, Bonfield W. Application of long chain amine activator in conventional acrylic bone cement. J Biomed Mater Res 1998; 43: 131-139. https://doi.org/10.1002/(sici)1097-4636(199822)43:2%3C131::aid-jbm7%3E3.0.co;2-p.

- de la Torre B, Fernández M, Vázquez B, Collía F, de Pedro JA, López-Bravo A, Román JS. Biocompatibility and other properties of acrylic bone cements prepared with antiseptic activators. J Biomed Mater Res B Appl Biomater 2003; 66: 502-513. https://doi.org/10.1002/jbm.b.10013.

- Ormsby R, McNally T, Mitchell C, Dunne N. Incorporation of multiwalled carbon nanotubes to acrylic based bone cements: effects on mechanical and thermal properties. J Mech Behav Biomed Mater 2010; 3: 136-145. https://doi.org/10.1016/j.jmbbm.2009.10.002.

- Khaled SMZ, Charpentier PA, Rizkalla. Physical and mechanical properties of PMMA bone cement reinforced with nano-sized titania fibers. Journal of Biomaterials Applications 2011:25:515-537. https://doi.org/10.1177/0885328209356944.

- Zhu W, Liu F, He J. Effect of polymerizable quaternary ammonium monomer MEIM-x's alkyl chain length and content on bone cement's antibacterial activity and physicochemical properties. Journal of the Mechanical Behavior of Biomedical Materials 2018; 87: 279-287. https://doi.org/10.1016/j.jmbbm.2018.08.004.

- Han, L-Y, Yang D, Wang H-L, Gao Z, Qi Y-Y, Zhou T, Xu Y, Chu J-J. Bone cement prepared using the isobornyl acrylate monomer shows decent mechanical strength and antibacterial activity. New Journal of Chemistry 2025;  49: 2106-2115. https://doi.org/10.1039/D4NJ04291C.

- Kim MJ, Park SY, Kang S, Lee YJ, Lee S-J, Kim JT. Song SY, Kim D-H, Oh SH. Low modulus PMMA-based bone cement for the reduction of vertebral fractures after vertebroplasty. Acta Biomaterialia 2025;203:399-411.  https://doi.org/10.1016/j.actbio.2025.07.053.

- Santos JGF, Pita VJRR, Melo PA, Nele M, Pinto JC. Effect of process variables on the preparation of artificial bone cements. Brazilian Journal of Chemical Engineering 2013; 30, 865-876. https://dx.doi.org/10.1590/S0104-66322013000400018.

- Tang Y, Chen L, Wu Z, Zhao K, Tan Q. Fabrication of injectable and expandable PMMA/PAASf bone cements. Composites Science and Technology 2017; 146: 203-209. https://doi.org/10.1016/j.compscitech.2017.04.032.

- Cui X, Huang C, Zhang M, Ruan C, Peng S, Li L, et al. Enhanced osteointegration of poly(methylmethacrylate) bone cements by incorporating strontium-containing borate bioactive glass. Journal of the Royal Society Interface 2017; 14: https://doi.org/10.1098/rsif.2016.1057.

- Bakhtiari SSE, Karbasi S, Hassanzadeh Tabrizi SA, Ebrahimi-Kahrizsangi R, Salehi H. Evaluation of the effects of chitosan/multiwalled carbon nanotubes composite on physical, mechanical and biological properties of polymethyl methacrylate-based bone cements. Materials Technology 2020; 35: 267-280. https://doi.org/10.1080/10667857.2019.1678086.

- Chuan D, Zhang L, Leng C, Chen Q, Miyazaki T, Liu J. Setting behavior, apatite-forming ability, mechanical strength of polymethylmethacrylate bone cement through bioactivity modification of phosphate functional groups combined with Ca2+ ions. J Biomater Sci Polym Ed 2020; 31: 2128-2151. https://doi.org/10.1080/09205063.2020.1795459.

- Hamdy TM. Modification of polymethylmethacrylate bone cement with halloysite clay nanotubes. BMC Oral Health 2024; 24:893. https://doi.org/10.1186/s12903-024-04600-3.

- Simorangkir APN , Kurniawati F , Prajatelistia E, Arjasa OP. The synthesis of polymethylmethacrylate (PMMA)                by miniemulsion polymerization for bone cement application: the role of anionic surfactant and co-stabilizer virgin coconut oil concentration. BIO Web of Conferences 2024; 127: 02002. https://doi.org/10.1051/bioconf/202412702002.

- Tabatabaee S, Delyanee M, Samanipour R, Marzban A, Tavakoli A, Alizadeh A. Advances in bone repair:                            a bioengineered PMMA-Based bone cement with enhanced otseoconductivity utilizing human-derived bone powder.     J Polym Environ 2025; 33, 2991-3012 (2025). https://doi.org/10.1007/s10924-025-03573-w.

- Ginebra MP, Gil FX, Planell JA, et al. Relationship between the morphology of PMMA particles and properties of acrylic bone cements. J Mater Sci: Mater Med 1996; 7: 375-379. https://doi.org/10.1007/BF00154553.

- Milner R. The development of theoretical relationships between some handling parameters (setting time and setting temperature), composition (relative amounts of initiator and activator) and ambient temperature for acrylic bone cement. J Biomed Mater Res B Appl Biomater 2004; 68: 180-185. https://doi.org/10.1002/jbm.b.20019.

- Przesławski G, Szcześniak K, Gajewski P, Marcinkowska A. Influence of initiator concentration on the polymerization course of methacrylate bone cement. Polymers 2022; 14: 5005. https://doi.org/10.3390/polym14225005.

- Vazquez B, Deb S, Bonfield W. Optimization of benzoyl peroxide concentration in an experimental bone cement based on poly(methyl methacrylate). J Mater Sci Mater Med 1997; 8: 455-460. https://doi.org/10.1023/a:1018565807727.

- Pascual B, Gurruchaga M, Ginebra MP, et al. Modified acrylic bone cement with high amounts of ethoxytriethyleneglycol methacrylate. Biomaterials 1999; 20: 453-463. https://doi.org/10.1016/s0142-9612(98)00191-4.

- Girija M, Kumar TS. Characterization of akermanite (AKT) and zirconia-infused PMMA bone cement composite with   superior physicochemical, mechanical, and bioactive properties for enhanced orthopaedic performance.       Composite Interfaces 2025; 32: 817-841. https://doi.org/10.1080/09276440.2024.2442169.

- Vazquez B, Deb S, Bonfield W, San Roman J. Characterization of new acrylic bone cements prepared with oleic acid derivatives. Journal of Biomedical Materials Research (Applied Biomaterials) 2002;63:88-97. https://doi.org/10.1002/jbm.10092.

- Madigan S, Towler MR, Lewis G. Influence of two changes in the composition of an acrylic bone cement on some of its properties: the case of Surgical Simplex® P.     J Mater Sci 2006; 41: 5758-5759. https://doi.org/10.1007/s10853-006-0107-7.

- Madigan S, Towler MR, Lewis G. Optimisation of the composition of an acrylic bone cement: application to relative amounts of the initiator and the activator/co-initiator in Surgical Simplex P. J Mater Sci Mater Med 2006; 17:                   307-311. https://doi.org/10.1007/s10856-006-8227-9.

- Moritz AR, Henriques FC. Studies of thermal injury- The relative importance of time and surface temperature in the  causation of cutaneous burns. Am J Pathol 1947; 23: 695-720. PMID: 19970955; PMCID: PMC1934304.

- Feith R. Side-effects of acrylic cement implanted into bone: a histological, (micro)angiographic, fluorescence-microscopic and autoradiographic study in the rabbit femur. Acta Orthopaedica Scandinavica 1975; 46: 1-136. https://doi.org/10.3109/ort.1975.46.suppl-161.01.

- Dickman R, Hall CK. Equation of state for chain molecules: Continuous-space analog of Flory theory.                                 The Journal of Chemical Physics 1986; 85: 4108-4115. https://doi.org/10.1063/1.450881.

- Eriksson AR, Albrektsson T. Temperature threshold levels for heat-induced bone tissue injury: a vital-microscopic study in the rabbit. J Prosthet Dent. 1983; 50: 101-107. https://doi.org/10.1016/0022-3913(83)90174-9.

- Pastor YY, Orellana J, Calle F, Pastor JY. Influence of wt. % and specific area of HRGO nanoplatelets addition on physical, thermal, mechanical behaviour and processing evolution of PMMA bone cement for orthopaedic and trauma surgery. 2021. http://dx.doi.org/10.20944/preprints202105.0213.v1.

- Krishnan MR, Alshabib RM, Alsharaeh EH. Effect of silver/reduced graphene oxide@titanium dioxide (Ag/rGO@TiO2) nanocomposites on the mechanical characteristics and biocompatibility of poly(styrene-co-methyl methacrylate)-based bone cement. Polymers (Basel) 2025; 17: 1970. https://doi.org/10.3390/polym17141970.

- Ayora-Gutiérrez G, Canché-Escamilla G, Cervantes-Uc JM, Uribe-Calderon JA. PMMA-bead decorated graphene oxide, a novel strategy to improve GO dispersion in bone cements. International Journal of Polymeric Materials and Polymeric Biomaterials 2025; 1-14. http://dx.doi.org/10.1080/00914037.2025.2516635.

- Gonçalves G, Cruz SM, Ramalho A, Grácio J, Marques PA. Graphene oxide versus functionalized carbon nanotubes as a reinforcing agent in a PMMA/HA bone cement.  Nanoscale 2012;4:2937-2945. https://doi.org/10.1039/c2nr30303e.

- Humez M, Fröschen FS, Wirtz DC, Kühn KD. Moderne Zementiertechnik der dritten Generation in der Knie-und Hüftendoprothetik [The third-generation modern cementing technique in hip and knee arthroplasty].                Orthopadie (Heidelb) 2023; 52: 968-980. https://doi.org/10.1007/s00132-023-04446-7.

- Bu WH, Asilebieke A, Han LY, Xu Y, Zhou T, Chu JJ. Preparation and characterization of novel PMMA bone cement containing 3,4-dichloro-5-hydroxyfuran-2(5H)-one. RSC Advances 2025;15:806-813. https://doi.org/10.1039/d4ra06607c.

- Nelson SB, Pinkney JA, Chen AF, Tande AJ. periprosthetic joint infection: current clinical challenges. Clin Infect Dis 2023; 77: e34-e45. https://doi.org/10.1093/cid/ciad360.

- Zhou H, Yang Y, Zhang Y, et al. Current status and perspectives of diagnosis and treatment of periprosthetic joint infection. Infect Drug Resist 2024; 17: 2417-2429. https://doi.org/10.2147/idr.s457644.

- Chong YY, Lau CML, Jiang T, Wen C, Zhang J, Cheung A, Luk MHL, Leung KCT, Cheung MH, Fu H, Chiu Y,            Chan PK.  Predicting periprosthetic joint infection in primary total knee arthroplasty: a machine learning model integrating preoperative and perioperative risk factors. BMC Musculoskeletal Disorders 2025;26:241. 2025;26:241.
https://doi.org/10.1186/s12891-025-08296-6.

- Cao Tang P, Chai H, Ma W, Lin B, Zhu Y, Abdirahman A, Xiao W, Zhang J, Li Y, Liu S, Wen T.  The application of antibiotic-loaded bone cement in preventing periprosthetic joint infection: an umbrella review. Journal of Orthopaedics and Traumatology 2025;26:23. https://doi.org/10.1186/s10195-025-00839-w

- Magruder ML, Heckmann ND, Lieberman JR, Scuderi G, Lustig S, Parvizi J, Mont MA. Novel technologies in periprosthetic joint infections: emerging therapeutics. J Arthroplasty 2025; in press.

97. Semeshchenko D, Slullitel PA,  Farinati A,  Albani-Forneris AF,  Piuzzi NS, Buttaro MA. Unconventional therapies in

periprosthetic joint infections: prevention and treatment: a narrative review. Journal of Clinical Medicine 2005; 14:

2610. https://doi.org/10.3390/jcm14082610.

- Lewis G. Antibiotic-free antimicrobial poly (methyl methacrylate) bone cements: A state-of-the-art review.                  World J Orthop 2022; 13: 339-353. https://doi.org/10.5312/wjo.v13.i4.339.

- DeBiase JM. A clinical review of prosthetic joint infections. Med Clin North Am 2025; 109: 615-623.

https://doi.org/10.1016/j.mcna.2024.12.003.

- Lin H, Gao Z, Shan T. et al. A review on the promising antibacterial agents in bone cement–

From past to current insights. J Orthop Surg Res 2024; 19: 673. https://doi.org/10.1186/s13018-024-05143-7.

- Yang L, Yergeshov AA, Al-Thaher Y, et al. Nanocomposite orthopaedic bone cement

combining long-acting dual antimicrobial drugs. Biomaterials Advances 2023; 153: 213538.

https://doi.org/10.1016/j.bioadv.2023.213538.

- Xu Y, Lin H, Gao Z, et al. Injectable isoniazid-loaded bone cement based on hydrazone bonds achieving long-

term release and decent mechanical properties. J Mater Chem B 2024; 12: 4389-4397.

https://doi.org/10.1039/d3tb02661b.

- Aksu D, Büyük Nİ, Aksu BÇ, Meriç G, Köse GT. Comparison of different bone cement formulations containing

boron derivatives. J Biomed Mater Res B Appl Biomater 2025; 113: e35640. https://doi.org/10.1002/jbm.b.35640.

- Ullah I, Ju J, Song Y, Chen S, et al. PMMA bone cement with AgNP@CDs nanocomposite for infection control

and inflammation mitigation. Regenerative Biomaterials, 2025; 12. https://doi.org/10.1093/rb/rbaf086.

- Gao Z, Kan YC, Xie YH, Guo R, Li C, Asilebieke A, Xu Y, Chu JJ; A review on non-leaching antibacterial bone

cement for orthopedic surgery: From past to current insights. AIP Advances 2023; 13: 105034.

https://doi.org/10.1063/5.0174845.

- Wang X, Bu WH, Han LY, Han LX, et al. Furanone-based comonomer used to manufacture antibacterial bone

cement with simultaneously enhanced mechanical strength and antibacterial activity. Materials Advances 2025;

6: 6406-6415. https://doi.org/10.1039/D5MA00442J.

- Qi YY, Han LY, Han LX, Bu WH, Xu Y, Chu JJ. Formulation and characterization of a sortase A inhibitor-loaded

PMMA bone cement. RSC Adv 2025; 15: 33966-33975. https://doi.org/10.1039/d5ra04437e.

- Huang MH, Chen YH, Pu CC, et al. Entrapping imipenem into poly(methyl methacrylate) assemblies by

microfluidic process retains the activity against prosthetic joint infections following thermal treatment. Colloids

Surf B Biointerfaces 2025; 255: 114950. https://doi.org/10.1016/j.colsurfb.2025.114950.

- Powell, Blaskovich MAT, Hansford KA.  Targeted protein degradation: the new frontier of antimicrobial discovery?

ACS Infectious Diseases 2021;7:2050-2067. https://doi.org/10.1021/acsinfecdis.1c00203.

- Gundapaneni D, Goswami T. Thermal isotherms in PMMA and cell necrosis during total hip arthroplasty.

J Appl Biomater Funct Mater 2014; 12: 193-202. https://doi.org/10.5301/jabfm.5000196.