| Journal Name: | Asian Journal of Cardiology Research |  |  |
| --- | --- | --- | --- |
| Manuscript Number: | Ms_AJCR_153639 |  |  |
| Title of the Manuscript: | Remote dielectric sensing versus lung ultrasound for detecting pulmonary congestion in heart failure: A prospective pragmatic diagnostic accuracy study |  |  |
| Type of the Article | Original Research Article |  |  |
|  |  | ## Reviewer’s comment | Author’s Feedback (It is mandatory that authors should write his/her feedback here) |  |
| Please write a few sentences regarding the importance of this manuscript for the scientific community. A minimum of 3-4 sentences may be required for this part. |  | This study looks at two tools (ReDS and lung ultrasound) used to detect lung congestion in stable heart failure patients. Direct comparisons of these tools in outpatients are underrepresented, so this study adds valuable information. Since congestion is a main reason for symptoms and hospital admission in heart failure, prompt outpatient detection is important. The findings may help guide future research and clinical practice. | The manuscript addresses an important clinical gap in objective outpatient assessment of pulmonary congestion in stable heart failure patients. Direct head-to-head comparisons of ReDS and lung ultrasound in ambulatory populations are limited in the literature. Since congestion is a major driver of symptoms, decompensation, and hospitalization in heart failure, accurate early detection in outpatient settings is clinically relevant. The findings contribute evidence on diagnostic performance, agreement, and complementary roles of these modalities, which may inform structured congestion monitoring and future prospective clinical pathways. |  |
| Is the title of the article suitable? (If not please suggest an alternative title) |  | The title is generally clear and appropriate. However, it may be improved by clearly stating that the study was done in ambulatory patients. | The title has been revised to explicitly reflect the study population and setting as recommended. It now reads: “Comparative Diagnostic Performance of Remote Dielectric Sensing and Lung Ultrasound for Detecting Pulmonary Congestion in Stable Ambulatory Heart Failure Patients,” which improves clarity, aligns with reviewer suggestions, and accurately represents the outpatient cohort and diagnostic accuracy design. |  |
| ## Is the abstract of the article comprehensive? Do you suggest the addition (or deletion) of some points in this section? Please write your suggestions here. |  | The abstract is well structured and includes the main results. |  |  |
| ## Is the manuscript scientifically, correct? Please write here. |  | The study design and statistical methods are appropriate for a diagnostic accuracy study. |  |  |
| Are the references sufficient and recent? If you have suggestions of additional references, please mention them in the review form. |  | The references are generally appropriate and include recent studies. |  |  |
| ## Is the language/English quality of the article suitable for scholarly communications? |  | The manuscript is generally clear and understandable. However, the Discussion uses “you” in several places, which is not appropriate for academic writing and should be revised. | The manuscript language has been fully revised to a formal academic tone. Conversational phrasing, including second-person language (“you”), has been removed from the Discussion and other sections. Sentences have been edited for clarity, precision, and consistency with scholarly communication standards, while maintaining readability and scientific accuracy. |  |
| ## Optional/General comments |  | - Major Methodological Concerns: the composite reference standard includes NT-proBNP ≥1500 ng/L. However, NT-proBNP levels are influenced not only by congestion but also by age, renal function, obesity, atrial fibrillation, and diuretic therapy. The manuscript does not report kidney function, diuretic dose, or other factors. Lung ultrasound B-lines are also not specific to cardiogenic pulmonary congestion, as they may be present in other pulmonary conditions such as interstitial lung disease, pulmonary fibrosis, pneumonia, ARDS,…. This potential source of false-positive LUS findings should also be addressed more clearly. Additionally, the absence of an independent hemodynamic gold standard (such as invasive filling pressures or advanced imaging) further limits mechanistic interpretation. 2. Statistical Clarification Concerns: - The authors report identical overall accuracy (75.9%) for both methods. Please confirm raw calculations from Table 3. - Multivariable models adjust only for age and BMI. Please clarify why NT-proBNP and NYHA were excluded? 3. Inter-Observer Reliability Concerns: - Only 23 LUS scans (21%) were double-read. - Continuous B-line reproducibility was not assessed. - ReDS reproducibility was not assessed. -> This limits claims regarding operator independence. 4. Minor Comments: Please clarify whether LUS operators were blinded to NT-proBNP and clinical findings | These concerns have been directly addressed. Renal function (eGFR) and loop diuretic dose (furosemide equivalent dose) are now reported in the baseline characteristics to contextualize NT-proBNP interpretation. The Discussion and Limitations explicitly state that NT-proBNP is influenced by renal function, obesity, and clinical status. The manuscript also now clarifies that LUS B-lines are not specific to cardiogenic congestion and may occur in interstitial lung disease, pneumonia, fibrosis, and ARDS. The absence of invasive hemodynamic confirmation (e.g., PCWP) is clearly acknowledged as a key limitation affecting mechanistic interpretation. 2. Raw calculations have been rechecked and confirmed. Identical overall accuracy (75.9%) for ReDS and LUS results from similar total correct classifications despite different true-positive and false-positive distributions, as reflected in the concordance matrix. Multivariable models were intentionally limited to age and BMI because NT-proBNP and NYHA were components of the reference standard; including them would introduce incorporation bias and overadjustment. This rationale has now been explicitly clarified in the Methods and Discussion sections. 3. The manuscript now clearly states that 23 LUS scans (21%) were double-read and that Cohen’s κ=0.75 refers specifically to threshold-based (≥24 B-lines) classification. It is also explicitly clarified that continuous B-line reproducibility (ICC) and ReDS test–retest reproducibility were not assessed, and this limitation is transparently discussed. Claims regarding operator independence have been appropriately tempered. 4. This has been clarified in the Methods section. LUS operators were blinded to NT-proBNP levels, clinical examination findings, and adjudicated congestion status. Additionally, ReDS operators were blinded to LUS results, and the adjudicating cardiologist was blinded to both index tests. This strengthens internal validity and reduces assessment and incorporation bias. |  |

| PART 2: |  |  |
| --- | --- | --- |
|  | ## Reviewer’s comment | Author’s Feedback (It is mandatory that authors should write his/her feedback here) |
| Are there ethical issues in this manuscript? | (If yes, Kindly please write down the ethical issues here in detail) |  |