| Journal Name: | Journal of Advances in Medical and Pharmaceutical Sciences |
| --- | --- |
| Manuscript Number: | Ms_JAMPS_151465 |
| Title of the Manuscript: | A Quality by Design: Scientific Approach to Develop and Validate a Laser-Based Analytical Method for Particle Size Distribution of Dapagliflozin Propanediol |
| Type of the Article | Method Article |

| PART 1: Comments |  |  |
| --- | --- | --- |
|  | Reviewer’s comment Artificial Intelligence (AI) generated or assisted review comments are strictly prohibited during peer review. | Author’s Feedback (It is mandatory that authors should write his/her feedback here) No use of any AI generated tools (ChatGPT, COPILOT, etc.) |
| Please write a few sentences regarding the importance of this manuscript for the scientific community. A minimum of 3-4 sentences may be required for this part. | QbD is not a new idea; despite this, it is frequently underutilized in practice. This work provides critically needed literature examples of QbD to guide its utilization. | Agreed with reviewer comment |
| Is the title of the article suitable? (If not please suggest an alternative title) | Yes, I would update for grammar to “A Quality by Design Scientific Approach to Development and Validation of a Laser-Based Analytical Method for Particle Size Distribution of Dapagliflozin Propanediol”” | Agreed with reviewer comment and updated the Title |
| Is the abstract of the article comprehensive? Do you suggest the addition (or deletion) of some points in this section? Please write your suggestions here. | I suggest updating the first sentence to “Particle size, shape, and distribution of the active pharmaceutical ingredient (API) may significantly influence the physicochemical and biopharmaceutical performance of dosage forms.” | Agreed with reviewer comment and updated the sentence |
| Is the manuscript scientifically, correct? Please write here. | The general concepts of the paper are mostly accurate. However, there is some misalignment with the language used in the paper compared to ICH. For example, Dv10, Dv50,Dv50 are referred to as CQAs. This is not exactly correct. Dv10, Dv50, and Dv90s may be critical formulation attributes or critical material attributes that influence a critical quality attribute, such as blend uniformity (given that Dapagliflozin is orally delivered). Though this is highly dependent on the drug product itself, as Validation is always tied to the intended use. I believe the holistic concept of QbD is being missed in this work due to the incorrect usage of language. The analytical method will have critical method parameters that influence the method's accuracy, repeatability, etc. I would frame the language differently. It is not clear what the scientific rationale is for the descriptive statistics in Table 8. It appears the authors took the average and SD across different experiment settings. For example, the Mean value of 3.65 is the average of the three data points for Dv10 at stir speeds of 2000RPM,2500RPM,3000RPM. At first I thought the rationale was assessment of the effect of stir speed on the CMA, Dv10,50,and 90 but the text concludes that “a stirrer speed of 3000 rpm was selected as optimal to ensure uniform sample circulation, effective dispersion, and reproducible particle size measurements.” It is unclear how the data from Table 8 concludes this. In addition, there are no repetitions of measurements of the samples' experimental setting. Similar comments for Table 10 | We agree with the reviewer that referring to Dv10, Dv50, and Dv90 directly as Critical Quality Attributes (CQAs) was not fully aligned with ICH Q8/Q9/Q10 terminology. As correctly pointed out, these particle size descriptors are more appropriately described as critical material attributes (CMAs) or critical formulation attributes, which may influence downstream drug product CQAs (e.g., blend uniformity, dissolution performance, and bioavailability), particularly for an orally administered drug such as Dapagliflozin. The relevance of these attributes is indeed dependent on the intended use and the specific drug product context. Hence updated CQA section. The calculated mean, standard deviation, and %RSD values reflect variability across different experimental conditions and are therefore not indicative of method repeatability. These statistics were used only to illustrate the sensitivity of particle size to method variable parameters (Stirrer speed and sonication time). Based on the stabilization of particle size parameters, a sonication time of 3 minutes was selected as optimal for subsequent studies. Method precision and reproducibility were evaluated separately using replicate measurements under fixed experimental conditions. |
| Are the references sufficient and recent? If you have suggestions of additional references, please mention them in the review form. | Yes | Agreed with reviewer comment |
| Is the language/English quality of the article suitable for scholarly communications? | Yes | Agreed with reviewer comment |
| Optional/General comments | Stir Time appears inappropriately labeled on Figure 2, making it difficult to interpret. | Figure 2 only represent impact on particle size with increasing stirrer speed and sonication time up to an optimum region. Stirrer speed should be maintained throughout the measurement time. Hence Stirrer speed is important for study |

| PART 2: |  |  |
| --- | --- | --- |
|  | Reviewer’s comment | Author’s Feedback (It is mandatory that authors should write his/her feedback here) |
| Are there ethical issues in this manuscript? | No. |  |