Review Form 3

	

	Journal Name:
	Asian Oncology Research Journal 

	Manuscript Number:
	Ms_AORJ_155418

	Title of the Manuscript: 
	INSTITUTION-SPECIFIC VALIDATION OF CTV-PTV MARGINS IN HEAD AND NECK RADIOTHERAPY USING DAILY CONE BEAM COMPUTED TOMOGRAPHY ON HALCYON ELITE: A POPULATION - BASED STUDY

	Type of the Article
	Clinical Practice article


	PART  1: Comments



	
	Reviewer’s comment


	Author’s Feedback (It is mandatory that authors should write his/her feedback here)



	Please write a few sentences regarding the importance of this manuscript for the scientific community. A minimum of 3-4 sentences may be required for this part.


	
	

	Is the title of the article suitable?

(If not please suggest an alternative title)


	YES
	

	Is the abstract of the article comprehensive? Do you suggest the addition (or deletion) of some points in this section? Please write your suggestions here.


	YES
	

	Is the manuscript scientifically, correct? Please write here.
	YES
	

	Are the references sufficient and recent? If you have suggestions of additional references, please mention them in the review form.
	YES
	

	Is the language/English quality of the article suitable for scholarly communications?


	
	

	Optional/General comments


	· The van Herk margin recipe used pertains to uncorrected inter-fraction uncertainties; with daily online corrections, residual post-correction errors should be analyzed instead. Using pre-correction shifts likely overestimates setup uncertainty for margin derivation under daily IGRT.

· Rotational errors were not measured, and the couch lacks 6-DOF correction, leaving unquantified residuals that may be clinically relevant for H&N.

· An apparent unit/logic inconsistency: the lateral margin of 0.34 cm (3.4 mm) exceeds the stated institutional 3 mm margin, yet the paper repeatedly claims all computed margins are smaller than 3 mm.
· The number of fractions per patient, total number of CBCTs analyzed, and per-axis distributions (e.g., histograms, outliers) are not reported undermining reproducibility and statistical confidence.

· No assessment of residual errors after correction (e.g., post-shift verification, second CBCT) and no intrafraction motion analysis; both are key for modern IGRT margin determination.

· Exclusive reliance on bony registration without analyzing soft-tissue/target surrogates may bias errors for subsites where bony anatomy is a poor proxy for CTV motion.

· Cohort and data completeness: Report fraction counts per patient, total CBCTs, and fraction-level distributions per axis. Include summary statistics (e.g., medians, IQRs), outlier analyses, and temporal trends (e.g., late-course weight loss effects).

· Residual and intrafraction measurements: Consider a post-correction verification CBCT (spotlight or half-arc) on a subset of fractions and intrafraction monitoring (e.g., post-beam CBCT, surface imaging) to bound remaining uncertainties in a margin recipe appropriate for daily IGRT.

· Site/subsite stratification: Stratify by subsite or immobilization differences (e.g., shoulder positioning) to explain the larger lateral variability and derive subsite-specific recommendations if warranted.

· nstitution-specific margin validation is valuable for QA and continuous improvement. However, the principal quantitative takeaway here (lateral margin > 3 mm) contradicts the paper’s conclusion. Correcting this interpretation would sharpen the clinical message—potentially motivating anisotropic margins (e.g., 3 mm V/LG, 4 mm lateral) or protocol refinements (e.g., shoulder support) while further evaluating residual and intrafraction uncertainties.

· The work could have greater impact by:

·  (i) quantifying residual errors under the actual correction protocol; 

· (ii) providing confidence intervals and sensitivity analyses;

·  (iii) stratifying by subsite; and (iv) tying results to adaptive imaging and planning strategies that are increasingly standard.
	


	PART  2: 



	
	Reviewer’s comment
	Author’s Feedback (It is mandatory that authors should write his/her feedback here)



	Are there ethical issues in this manuscript? 


	(If yes, Kindly please write down the ethical issues here in detail)


	


Reviewer details:

Haider Mohammed Majeed, University of Baghdad – College of Nursing, Iraq

Created by: DR
              Checked by: PM                                           Approved by: MBM
   
Version: 3 (07-07-2024)

